| Company Name: | Biomate Medical Devices Technology Co., Ltd. |  |
| --- | --- | --- |
| company address: | 1F, No. 59, Luke 2nd Rd., Luzhu Dist., 821 Kaohsiung City, TAIWAN |  |
| SRN : | TW-MF-000002521 |  |
| Contact name: | Su Yu-Jung |  |
| Contact number: | +886-7-952-5999 #9888 |  |
| Year / Time Period for PSUR: | 1 year |  |
| Device family name: | Biomate Implant Systems |  |
| Product Type | Implant, Abutment, Screw, Surgical Instruments |  |
| Product classification level: | IIb, IIa |  |
| Technical file number: | BI-TD-0001, BI-TD-0002, BI-TD-0003 |  |
| Basic UDI-DI No.: | 471987540Implant5K, 471987540AbutmentVC, 471987540ScrewIIbUN, 471987540Surgical(IIa)FP |  |
| Date the PSUR was submitted to the Notified Body: | NA |  |
| Submitted by Eudamed (Class III or Implantable Devices [MDR] | NO ☐ YES Date：NA |  |
| Notified body name | UDEM ADRIATIC D.O.O |  |
| Notified body ID number | 2696 |  |
| Approver (Name/Title) | Signature | Date |
| Su Yu-Jung / PRRC |  | 2026.01.20 |

Revision Status

| Revision # | Date (DD/MM/YYYY) | Revision History |
| --- | --- | --- |
| V1.0 | 17/10/2022 | Initial release |
| V1.1 | 28/11/2022 | Revised C1, C2, C3 Chapter 2, add D, E, F, G |
| V1.2 | 01/09/2023 | annual update |
| V1.3 | 01/09/2024 | annual update |
| V1.4 | 08/04/2025 | Update clinical data and Sales data |
| V1.5 | 03/09/2025 | Update clinical data and Sales data |
| V1.6 | 20/01/2025 | Update Section 1, Secotion 5.4 |

1.	Device Description and Administrative Information	5

1.1.	Device Schedule and Status (Categorization of Devices)	5

1.2.	Justification for Grouping and Selection of the "Leading Device"	5

1.3.	Indications and Contraindications	6

2.	Conclusions of the benefit-risk determination:	7

3.	Main findings of the PMCF report:	9

A1 – Gathering of clinical experience gained:	9

A1.1.	Summarize all collected clinical data obtained from this activity:	9

A1.2.	List and justify any deviations from the plan:	33

A1.3.	Analysis of the findings (whether positive or negative):	33

A1.4.	Summarize the collected data:	37

A2 – Screening of scientific literature:	38

A2.1.	Summarize all collected clinical data obtained from this activity:	38

A2.2.	List and justify any deviations from the plan:	42

A2.3.	Analysis of the findings (whether positive or negative):	42

A2.4.	Summarize the collected data:	62

A3 – PMS publicly listed available databases:	62

A3.1.	Summarize all collected clinical data obtained from this activity:	62

A3.2.	List and justify any deviations from the plan:	75

A3.3.	Analysis of the findings (whether positive or negative):	75

A3.4.	Summarize the collected data:	77

B – Evaluation of clinical data relating to similar devices:	78

C – Impact of the results on the technical documentation:	78

D – Reference to any applicable common specification(s), harmonized standard(s) or guidance document(s) applied:	79

4.	Statistical data of unqualified quality	79

4.1.	Unqualified quality data	79

4.2.	Summarize the collected data of unqualified quality:	80

5.	Sales and Complaint Data	81

5.1.	Sales data	81

5.2.	Complaint Data	82

5.3.	Rational and description of any preventive and corrective actions taken:	82

5.4.	Summarize the collected data of the complaint:	83

6.	Reportable Events Trending	87

7.	Conclusions of the PSUR	89

This PSUR contains information from our PMS Plan activities and PMCF Plan for Biomate Implant Systems for the time period specified above.

- Device Description and Administrative Information

  - Device Schedule and Status (Categorization of Devices)

The following table lists all devices covered by this PSUR, including their administrative milestones, EMDN codes, and current market status, in accordance with MDCG 2022-21.

Table 1-1. Device Schedule, History, and Status

| Device Family | Risk Class | Basic UDI-DI | EMDN Code | Date of First DoC (YYYY/MM/DD) | Current Status |
| --- | --- | --- | --- | --- | --- |
| Implant | IIb | 471987540Implant5K | P01020101 | 2017/10/13 | Marketed |
| Abutment | IIb | 471987540AbutmentVC | P01020180 | 2017/10/13 | Marketed |
| Screw (IIb) | IIb | 471987540ScrewIIbUN | P01020180 | 2017/10/13 | Marketed |
| Surgical Instruments | IIa | 471987540Surgical(IIa)FP | Z12110102 | 2017/10/13 | Marketed |

Note 1: EMDN codes and Risk Classes are based on the Declaration of Conformity (Doc #: BI-IM-S001 Rev V1.2).

Note 2: The "Date of First DoC" refers to the initial release date (Rev. A) of the declaration.

  - Justification for Grouping and Selection of the "Leading Device"

This PSUR covers a group of devices belonging to the Biomate Implant Systems. To facilitate a comprehensive safety assessment, the Bone Type Implant (Basic UDI-DI: 471987540Implant5K) has been selected as the "Leading Device".

Rationale for Leading Device Selection:

- Highest Risk Classification: The Implant is classified as Class IIb (Rule 8), representing the highest risk classification within the device group.

- Clinical Relevance: As an implantable device in direct contact with bone tissue, it presents the most critical safety profile compared to accessories (abutments, screws) or surgical instruments.

- Data Coverage: The safety data and clinical performance of the entire system are fundamentally driven by the success and survival rate of the Implant. Therefore, the PSUR schedule and data collection period are driven by this Leading Device.

  - Indications and Contraindications

- Indications for Use: The Biomate Implant Systems are designed for the restoration of lost teeth. The Implant and the Screw are placed into the maxillary or mandibular alveolar bone by surgical procedures. The lost tooth will be rehabilitated after osseointegration of the Implant, connection of the Abutment to the Implant and Screw, and attachment to the prosthesis.

- Contraindications: The device is contraindicated for patients with active infection, inflammation, uncontrolled systemic diseases (e.g., diabetes), or insufficient bone volume/quality that compromises initial stability. (Please refer to the IFU for detailed contraindications).

- Conclusions of the benefit-risk determination:

| Based on the risk assessment results, 27 risks are near the risk boundary, while the remaining risk factors are within an acceptable range, as detailed in the table. The following provides a more detailed explanation: - Manufacturer's Risk Control: Risks numbered No. 1-7, 11-18, 21-23, and 27-37 have been further mitigated through design and internal procedural controls, reducing their likelihood of occurrence. These control measures include strict process monitoring and quality inspections to ensure that the product meets standards. The benefits to human health outweigh these risks, and therefore these risks are considered acceptable. - Physician's Risk Control: Risks numbered No. 8-10, 14, 19-21, 24-26, 38-40, 44-45, 47-48, 50, 53, 57, and 65 are potential risks arising from the physician's operations. Since medical devices must be operated by certified dentists, and patients should receive treatment in legitimate dental clinics, these risks are deemed to be effectively reduced through proper training and knowledge dissemination. Both physicians and patients are fully aware of the relevant treatment risks. | Risk distribution |
| --- | --- |
| The risk management measures have been included in the product instructions and are provided with the packaging. These labels help to prevent and reduce residual risks, and thus, these risks are considered acceptable. - Known Implant Surgery Risks: Implant surgeries involve known risks such as local swelling, pain, inflammation, and infection. Most of these risks can be managed by the physician with appropriate pre- and post-operative care. However, insufficient osseointegration can lead to implant failure, with causes such as poor bone quality, improper implantation process, poor oral hygiene, or premature loading. All implant systems carry the risk of needing re-implantation or removal due to inadequate osseointegration. Patients should be advised to regularly consult a professional dentist to detect and manage potential risks early. These risks are clearly marked in the product instructions, and it is anticipated that the risk can be effectively reduced, thus safeguarding human health. Therefore, these risks are considered acceptable. - Feedback and Improvement on the 2024 Screw Fracture Incident: In 2024, a complaint was received regarding a screw fracture. Upon investigation, it was found that the doctor selected an inappropriate specification, resulting in excessive pressure on the screw, which led to its fracture. To prevent similar incidents, it is recommended to provide doctors with appropriate specification guidelines. - Storage and Transportation Risk Control: In 2024, no complaints were received regarding product damage caused by high temperatures, low temperatures, impacts, or drops during storage or transportation. This indicates that the risk management in these areas is effective. - Review and Update of Clinical Literature: The identified literature has been thoroughly evaluated according to established standards during the literature search. These documents provide scientific evidence regarding the safety and performance of the Biomate implant system. Literature collection covers the years 1983 to 2023 and does not typically require annual updates due to the relatively slow pace of change in dental technology. A strategy of updating the literature every five years not only reduces the annual workload but also maintains the quality and integrity of the evaluation. However, if significant technological changes or new clinical evidence emerge, the literature should be updated in advance to ensure timeliness and accuracy of the evaluation. Based on the risk analysis and control measures implemented, 27 risks have been identified as being near the risk boundary, while all other risk factors fall within acceptable limits, as detailed in the table. The manufacturer has successfully reduced the frequency and severity of these risks through design and internal procedural risk controls, as well as mitigation measures implemented by qualified healthcare professionals. In this assessment, the benefits of using the product for maintaining human health outweigh the identified residual risks. Additionally, known implant surgery risks, such as bone loss or implant failure due to insufficient osseointegration, are manageable with appropriate pre- and post-operative care. Therefore, as long as users, including healthcare professionals and patients, follow the provided instructions and precautions, all residual risks are considered acceptable. To ensure ongoing safety and effectiveness, regular monitoring is recommended, and risk management strategies should be updated in response to new clinical evidence or technological advancements. |  |

- Main findings of the PMCF report:

## A1 – Gathering of clinical experience gained:

- Summarize all collected clinical data obtained from this activity:

The following table shows the CT taken at the patient's return visit.

Source from:

1. Prof. Dr. Chiang, Tao

The gathers were conducted from 2016/03/07 to 2024/09/1.

| Cases | SEX, Age, Country | Upper Jaw | Lower Jaw | Implant Specification | Follow-up Duration after Surgery | Follow-up (M.D) |
| --- | --- | --- | --- | --- | --- | --- |
| 1-L464 | male, 59 year old, Taiwan | (#16) (#26) | - | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2021.12.15~ 2022.04.27 | 4.12 |
|  |  |  |  | (#26) : Bone Type Implant Plus: Ø 4.5 × L10 | 2021.04.19~ 2022.02.16 | 9.28 |
| 1-R172 | Female, 48 year old, Canada | (#16) | - | (#16) : Bone Type Implant Plus: Ø 4.5 × L8 | 2021.11.24~ 2022.05.20 | 5.26 |
| 1-T174 | Female, 51 year old, Taiwan | (#15) (#16) | - | (#15) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.11.18~ 2022.05.18 | 6.0 |
|  |  |  |  | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2021.11.18~ 2022.05.18 | 6.0 |
| 1-F919 | male, 58 year old, Taiwan | (#25) (#26) | - | (#25) : Bone Type Implant: Ø 4.8 × L10 | 2021.09.28~ 2022.02.14 | 4.17 |
|  |  |  |  | (#26) : Bone Type Implant: Ø 4.8 × L8 | 2021.09.28~ 2022.02.14 | 4.17 |
| 1-D900 | male, 72 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 5.5 × L10 | 2021.09.01~ 2022.01.26 | 4.25 |
| 1-T280 | male, 57 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant Plus: Ø 4.5 × L8 | 2021.08.19~ 2021.12.06 | 3.17 |
| 1-S408 | Female, 27 year old, Taiwan | (#16) (#26) | - | (#16) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.07.30~ 2022.02.14 | 6.15 |
|  |  |  |  | (#26) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.07.30~ 2022.02.14 | 6.15 |
| 1-F668 | Female, 59 year old, Taiwan | (#26) (#27) | - | (#26) : Bone Type Implant: Ø 4.8 × L8 | 2021.06.04~ 2022.11.19 | 17.15 |
|  |  |  |  | (#27) : Bone Type Implant Plus: Ø 4.5 × L10 | 2021.06.04~ 2022.11.19 | 17.15 |
| 1-S603 | male, 41 year old, England | (#26) | - | (#26) : Bone Type Implant: Ø 5.5 × L10 | 2021.04.07~ 2022.04.25 | 12.18 |
| 1-Q847 | male, 56 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 5.5 × L10 | 2021.03.30~ 2021.08.04 | 4.5 |
| 1-O493 | Female, 57 year old, Taiwan | (#15) | - | (#15) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.03.22~ 2021.08.02 | 4.11 |
| 1-S849 | Female, 55 year old, Taiwan | (#14) (#15) | - | (#14) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.03.18~ 2021.11.16 | 7.29 |
|  |  |  |  | (#15) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.03.18~ 2021.11.16 | 7.29 |
| 1-I474 | Female, 51 year old, Taiwan | (#25) | - | (#25) : Bone Type Implant Plus: Ø 4.0 × L10 | 2021.03.11~ 2021.08.01 | 4.21 |
| 1-B708 | Female, 66 year old, Taiwan | (#17) | - | (#17) : Bone Type Implant Plus: Ø 4.5 × L10 | 2020.11.10~ 2021.02.25 | 3.15 |
| 1-H114 | male, 62 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant Plus: Ø 4.5 × L10 | 2020.11.02~ 2021.04.01 | 4.30 |
| 1-S893 | Female, 42 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 4.8 × L10 | 2020.06.19~ 2020.11.11 | 4.23 |
| 1-S458 | Female, 54 year old, Taiwan | (#16) (#27) | - | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2020.12.07~ 2021.04.14 | 4.7 |
|  |  |  |  | (#27) : Bone Type Implant: Ø 5.5 × L10 | 2020.03.04~ 2020.07.01 | 3.27 |
| 1-F468 | male, 74 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2019.04.19~ 2019.09.09 | 4.21 |
| 1-P793 | male, 61 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2019.01.18~ 2019.06.04 | 4.17 |
| 1-K104 | male, 55 year old, Taiwan | (#15) (#25) | - | (#15) : Bone Type Implant: Ø 4.8 × L12 | 2019.03.11~ 2019.09.06 | 5.26 |
|  |  |  |  | (#25) : Bone Type Implant: Ø 5.5 × L10 | 2018.01.19~ 2019.03.11 | 13.2 |
| 1-S009 | Female, 94 year old, Taiwan | (#25) (#26) | - | (#25) : Bone Type Implant: Ø 5.5 × L10 | 2018.12.11~ 2019.06.18 | 6.7 |
|  |  |  |  | (#26) : Bone Type Implant: Ø 5.5 × L8 | 2018.12.11~ 2019.06.18 | 6.7 |
| 1-R595 | Female, 41 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 4.8 × L12 | 2018.09.20~ 2020.04.30 | 19.10 |
| 1-D536 | Female, 70 year old, Taiwan | (#25) | - | (#25) : Bone Type Implant Plus: Ø 4.0 × L10 | 2018.09.18~ 2019.03.05 | 5.15 |
| 1-E039 | male, 58 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 5.5 × L10 | 2018.02.26~ 2021.06.09 | 39.14 |
| 1-J590 | Female, 56 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 5.5 × L8 | 2017.11.03~ 2018.04.27 | 5.24 |
| 1-K304 | Female, 61 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 4.8 × L10 | 2017.10.12~ 2018.05.29 | 7.17 |
| 1-J868 | male, 67 year old, Taiwan | (#15) | - | (#15) : Bone Type Implant: Ø 4.1 × L10 | 2017.09.15~ 2018.07.17 | 10.2 |
| 1-L771 | male, 23 year old, Taiwan | (#25) | - | (#25) : Bone Type Implant: Ø 5.5 × L10 | 2017.05.04~ 2017.12.21 | 7.17 |
| 1-E093 | Female, 51 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 4.8 × L8 | 2017.01.13~ 2017.06.13 | 5.0 |
| 1-R127 | male, 74 year old, Taiwan | (#27) | - | (#27) : Bone Type Implant: Ø 5.5 × L10 | 2016.11.08~ 2017.07.10 | 8.2 |
| 1-R300 | Female, 58 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø 4.1 × L10 | 2016.10.13~ 2017.03.14 | 5.1 |
| 1-D261 | Female, 60 year old, Taiwan | (#17) | - | (#17) : Bone Type Implant: Ø 5.5 × L8 | 2016.06.21~ 2018.05.11 | 22.20 |
| 1-O899 | Female, 50 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 4.8 × L8 | 2016.05.18~ 2018.07.13 | 25.25 |
| 1-H794 | Female, 53 year old, Taiwan | (#16) (#17) | - | (#16) : Bone Type Implant: Ø 4.8 × L8 | 2016.03.07~ 2019.09.25 | 42.18 |
|  |  |  |  | (#17) : Bone Type Implant: Ø 4.8 × L8 | 2016.03.07~ 2019.09.25 | 42.18 |

2. Dr. FUYI,LIN

The gathers were conducted from 2021/10/01 to 2025/02/06

| Cases | SEX, Age, Country | Upper Jaw | Lower Jaw | Implant Specification | Follow-up Duration after Surgery | Follow-up Survival (M.D) |
| --- | --- | --- | --- | --- | --- | --- |
| 2-530426 | Female, 60 year old, Taiwan | - | (#46) (#47) | (#46) : Bone Type Implant:Ø4.8× L10 | 2023.08.25~ 2024.04.02 | 7.8 |
|  |  |  |  | (#47) : Bone Type Implant:Ø4.8× L10 | 2023.08.25~ 2024.04.02 |  |
| 2-681015 | male, 42 year old, Taiwan |  | (#46) (#47) | (#46) : Bone Type Implant:Ø4.1× L10 | 2021.10.01~ 2023.06.17 | 20.16 |
|  |  |  |  | (#47) : Bone Type Implant:Ø4.1× L10 | 2021.10.01~ 2023.06.17 |  |
| 2-500615 | male, 62 year old, Taiwan | (#21) (#22) (#23) (#24) | - | (#21) : Bone Type Implant:Ø3.3× L14 | 2022.03.10~ 2024.04.20 | 25.10 |
|  |  |  |  | (#22) : Bone Type Implant:Ø3.3× L14 |  |  |
|  |  |  |  | (#23) : Bone Type Implant:Ø4.1× L12 |  |  |
|  |  |  |  | (#24) : Bone Type Implant Plus:Ø3.5× L14 |  |  |
| 2-0721225 | male, 38 year old, Taiwan | (#16) | - | (#16): Bone Type Implant Plus:Ø4.5 × L10 | 2022.04.13~ 2022.11.14 | 7.1 |
| 2-460921 | Female, 65 year old, Taiwan | - | (#46) (#47) | (#46) : Bone Type Implant Plus:Ø4.0× L8 | 2022.04.12~ 2022.12.29 | 8.17 |
|  |  |  |  | (#47) : Bone Type Implant Plus:Ø4.0× L8 |  |  |
| 2-0540709 | male, 57 year old, Taiwan | (#16) | - | (#16) : Bone Type Implant: Ø 4.1 × L12 | 2022.04.27~ 2023.02.27 | 10.0 |
| 2-781020 | male, 32 year old, Taiwan | - | (#46) | (#46): Bone Type Implant:Ø4.1× L10 | 2022.04.28~ 2022.11.12 | 6.15 |
| 2-0480709 | male, 62 year old, Taiwan | - | (#34)(#35) (#37)(#45) (#47) | (#34) : Bone Type Implant: Ø 4.8 × L8 | 2022.05.23~ 2024.10.16 | 28.23 |
|  |  |  |  | (#35) : Bone Type Implant Plus: Ø4.5 × L8 |  |  |
|  |  |  |  | (#37) : Bone Type Implant Plus: Ø4.5 × L8 |  |  |
|  |  |  |  | (#45) : Bone Type Implant Plus: Ø4.5 × L8 |  |  |
|  |  |  |  | (#47) : Bone Type Implant Plus: Ø4.5 × L8 |  |  |
| 2-0650123 | male, 46 year old, Taiwan |  | (#46) | (#46) : Bone Type Implant: Ø 4.8 × L12 | 2022.05.25~ 2023.07.12 | 13.17 |
| 2-0501112 | Female, 61 year old, Taiwan | (#26) | - | (#26) : Bone Type Implant: Ø4.1 × L12 | 2023.05.10~ 2024.01.24 | 8.14 |
| 2-0480901 | male, 62 year old, Taiwan | (#17) | - | (#17) : Bone Type Implant: Ø4.8 × L10 | 2022.06.15~ 2023.12.27 | 18.12 |
| 2-0401111 | male, 70 year old, Taiwan | (#12) (#13) (#15) (#16) (#22) (#23) (#25) (#26) | - | (#12) : Bone Type Implant: Ø3.3 × L10 | 2022.06.13~ 2024.01.31 | 19.18 |
|  |  |  |  | (#13) : Bone Type Implant: Ø3.3 × L10 |  |  |
|  |  |  |  | (#15) : Bone Type Implant: Ø4.8 × L8 |  |  |
|  |  |  |  | (#16) : Bone Type Implant: Ø 4.8 × L8 |  |  |
|  |  |  |  | (#22) : Bone Type Implant: Ø 3.3 × L10 |  |  |
|  |  |  |  | (#23) : Bone Type Implant: Ø 3.3 × L10 |  |  |
|  |  |  |  | (#25) : Bone Type Implant: Ø4.8 × L8 |  |  |
|  |  |  |  | (#26) : Bone Type Implant: Ø4.8 × L8 |  |  |
|  |  | - | (#31) (#33) (#35) (#36) (#41) (#43) (#45) (#46) | (#31) : Bone Type Implant: Ø 3.3 × L10 | 2022.08.01~ 2024.01.31 | 17.30 |
|  |  |  |  | (#33) : Bone Type Implant: Ø3.3 × L10 |  |  |
|  |  |  |  | (#35) : Bone Type Implant: Ø4.1 × L10 |  |  |
|  |  |  |  | (#36) : Bone Type Implant: Ø 4.1 × L10 |  |  |
|  |  |  |  | (#41) : Bone Type Implant: Ø3.3 × L10 |  |  |
|  |  |  |  | (#43) : Bone Type Implant: Ø3.3 × L10 |  |  |
|  |  |  |  | (#45) : Bone Type Implant: Ø4.1 × L10 |  |  |
|  |  |  |  | (#46) : Bone Type Implant: Ø4.1 × L8 |  |  |
| 2-0600223 | male, 51 year old, Taiwan | (#25) | - | (#25) : Bone Type Implant: Ø 4.1 × L12 | 2022.06.22~ 2024.07.10 | 24.18 |
|  |  | - | (#35) (#37) | (#35) : Bone Type Implant: Ø4.1 × L12 | 2023.11.22~ 2024.07.10 | 7.18 |
|  |  |  |  | (#37) : Bone Type Implant: Ø 4.8× L12 |  |  |
| 2-0700802 | male, 42 year old, Taiwan | (#25) | - | (#25): Bone Type Implant:Ø4.8 × L10 | 2022.06.22~ 2024.01.31 | 19.9 |
| 2-0430621 | male, 68 year old, Taiwan | (#15) | - | (#15) : Bone Type Implant: Ø4.1 × L12 | 2022.06.22~ 2024.09.11 | 26.20 |
| 2-451106 | Female, 65 year old, Taiwan | (#24) | - | (#24): Bone Type Implant:Ø3.5 × L12 | 2022.06.21~ 2024.09.07 | 26.17 |
|  |  | - | (#32) (#41) (#43) | (#32): Bone Type Implant:Ø4.1 × L12 | 2024.05.23~ 2024.09.07 | 3.15 |
|  |  |  |  | (#41): Bone Type Implant:Ø4.1 × L12 |  |  |
|  |  |  |  | (#43): Bone Type Implant:Ø4.1 × L12 |  |  |
| 2-0531118 | male, 57 year old, Taiwan | (#16) | - | (#16): Bone Type Implant Plus: Ø5.5 × L10 | 2022.07.06~ 2024.07.10 | 24.4 |
|  |  | (#26) | - | (#26) : Bone Type Implant: Ø 4.8 × L10 | 2022.12.21~ 2024.07.10 | 18.19 |
| 2-540228 | Female, 58 year old, Taiwan | (#16) | - | (#16): Bone Type Implant:Ø4.8 × L10 | 2022.07.12~ 2023.04.22 | 9.10 |
| 2-500512 | Female, 61 year old, Taiwan | (#14) | - | (#14): Bone Type Implant:Ø3.3 × L10 | 2022.07.15~ 2024.05.16 | 22.1 |
|  |  | - | (#36) | (#36): Bone Type Implant:Ø4.1 × L10 | 2024.05.16~ |  |
| 2-520818 | Female, 58 year old, Taiwan | (#24)(#25) (#26) | - | (#24) : Bone Type Implant: Ø4.1 × L10 | 2022.08.16~ 2024.07.09 | 22.23 |
|  |  |  |  | (#25) : Bone Type Implant: Ø 4.1 × L10 |  |  |
|  |  |  |  | (#26) : Bone Type Implant: Ø4.8 × L10 |  |  |
|  |  | - | (#34)(#35) (#36) | (#34) : Bone Type Implant: Ø 4.1 × L10 | 2022.09.16~ 2024.07.09 | 21.23 |
|  |  |  |  | (#35) : Bone Type Implant: Ø 4.1 × L10 |  |  |
|  |  |  |  | (#36) : Bone Type Implant: Ø4.8 × L10 |  |  |
|  |  | - | (#46) | (#46) : Bone Type Implant: Ø 4.8 × L12 | 2024.07.09~ |  |
| 2-0550122 | Female, 56 year old, Taiwan | - | (#36) | (#36) : Bone Type Implant: Ø 4.8 × L12 | 2022.09.19~ 2024.06.24 | 21.5 |
|  |  | (#15) | - | (#15) : Bone Type Implant: Ø4.8 × L10 | 2023.08.09~ 2024.06.24 | 10.15 |
|  |  | - | (#37) | (#37) : Bone Type Implant: Ø 4.8 × L12 | 2024.01.17. |  |
|  |  | (#14) | - | (#14) : Bone Type Implant: Ø3.3 × L12 | 2024.06.24 |  |
| 2-0830806 | Female, 28 year old, Taiwan | (#16) | - | (#14) : Bone Type Implant: Ø 4.8 × L10 | 2022.09.19~ 2023.07.24 | 10.5 |
| 2-520204 | Female, 59 year old, Taiwan | - | (#35) (#36) | (#35) : Bone Type Implant: Ø4.1 × L8 | 2022.09.22~2 023.11.27 | 14.5 |
|  |  |  |  | (#36) : Bone Type Implant: Ø4.8 × L8 |  |  |
| 2-0520731 | Female, 59 year old, Taiwan | (#14) | - | (#14) : Bone Type Implant: Ø 4.1 × L12 | 2022.10.03~ 2023.05.08 | 7.5 |
|  |  | - | (#46) | (#46) : Bone Type Implant: Ø 4.8 × L10 |  |  |
| 2-560823 | Female, 55 year old, Taiwan | - | (#47) | (#47): Bone Type Implant:Ø 4.8× L12 | 2022.10.27~ 2023.11.17 | 12.21 |
| 2-761130 | Female, 35 year old, Taiwan | (#25) | - | (#25): Bone Type Implant:Ø4.1 × L10 | 2022.11.11~ 2023.07.07 | 7.26 |
| 2-0520226 | Female, 60 year old, Taiwan | (#16) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2022.11.28~ 2023.07.24 | 7.26 |
| 2-790220 | male, 33 year old, Taiwan | - | (#46) | (#46): Bone Type Implant Plus: Ø4.5× L14 | 2022.12.15~ 2023.06.02 | 5.18 |
| 2-0810722 | male, 31 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2022.12.26~ 2024.01.17 | 12.22 |
| 2-0330816 | male, 79 year old, Taiwan | (#16) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.02.01~ 2023.10.04 | 8.3 |
| 2-0571223 | Female, 54 year old, Taiwan | - | (#36) (#37) | (#36): Bone Type Implant Plus: Ø4.5 × L10 | 2023.02.01~ 2023.10.02 | 8.1 |
|  |  |  |  | (#37): Bone Type Implant Plus: Ø4.5× L10 |  |  |
| 2-530211 | male, 60 year old, Taiwan | - | (#37) | (#37): Bone Type Implant: Ø4.8× L10 | 2023.03.07~ 2024.12.11 | 21.4 |
| 2-0750925 | Female, 36 year old, Taiwan |  | (#37) | (#37): Bone Type Implant: Ø4.8× L12 (Immediate implant placement) | 2023.03.15~ 2023.10.11 | 6.26 |
| 2-0560810 | male, 55 year old, Taiwan |  | (#43) (#44) (#45) (#46) | (#43): Bone Type Implant: Ø4.1× L8 | 2023.03.15~ 2024.06.03 | 14.19 |
|  |  |  |  | (#44): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#45): Bone Type Implant: Ø4.1× L8 |  |  |
|  |  |  |  | (#46): Bone Type Implant: Ø4.8× L8 |  |  |
| 2-540302-1 | Female, 57 year old, Taiwan | (#14) (#16) (#24) (#26) | - | (#14): Bone Type Implant: Ø4.1× L10 | 2022.09.28~ 2023.12.26 | 14.28 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#24): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-520415-1 | Female, 59 year old, Taiwan | (#21) | - | (#21): Bone Type Implant: Ø4.1× L12 (Immediate implant placement) | 2023.03.28~ | 5.2 |
| 2-0670718 | Female, 44 year old, Taiwan | (#15) (#16) (#17) (#24) (#25) (#26) (#27) | - | (#15): Bone Type Implant: Ø3.3× L12 | 2023.03.29~ 2024.07.29 | 16 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#24): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#25): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.1× L12 |  |  |
| 2-0811223 | Female, 30 year old, Taiwan | (#15) (#16) | - | (#15): Bone Type Implant: Ø4.0× L12 | 2023.04.03~ 2023.10.16 | 6.13 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-571123 | male, 54 year old, Taiwan | (#11) (#12) | (#36) (#46) | (#11): Bone Type Implant: Ø4.1× L12 | 2024.06.04~ 2024.12.10 | 6.6 |
|  |  |  |  | (#12): Bone Type Implant: Ø3.3× L12 | 2024.06.04~ 2024.12.10 | 6.6 |
|  |  |  |  | (#36): Bone Type Implant: Ø4.1× L12 | 2023.04.18~ 2024.12.10 | 19.22 |
|  |  |  |  | (#46): Bone Type Implant: Ø4.8× L10 | 2024.05.16~ 2024.12.10 | 6.24 |
| 2-0440107 | Female, 68 year old, Taiwan | - | (#31) (#33) (#41) (#43) (#44) | (#31): Bone Type Implant: Ø4.1× L10 | 2023.04.24~ 2024.12.09 | 19.15 |
|  |  |  |  | (#33): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#41): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#43): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#44): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-630924 | Female, 48 year old, Taiwan | - | (#35) (#37) | (#35): Bone Type Implant: Ø4.1× L12 | 2023.04.28~ 2024.10.01 | 17.3 |
|  |  |  |  | (#37): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0431025 | Female, 68 year old, Taiwan | - | (#35) (#36) (#37) | (#35): Bone Type Implant: Ø3.3× L12 | 2023.05.22~ 2024.10.09 | 16.17 |
|  |  |  |  | (#36): Bone Type Implant: Ø4.8× L8 |  |  |
|  |  |  |  | (#37): Bone Type Implant: Ø4.8× L8 |  |  |
| 2-0581012 | male, 53 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.05.22~ 2024.01.10 | 7.19 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0550623 | Female, 56 year old, Taiwan | - | (#45) (#46) (#47) | (#45): Bone Type Implant: Ø4.1× L12 | 2022.09.14~ 2023.05.22 | 8.8 |
|  |  |  |  | (#46): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#47): Bone Type Implant: Ø4.8× L12 |  |  |
| 2-0401003 | Female, 56 year old, Taiwan | (#12) (#21) | - | (#12): Bone Type Implant: Ø3.3× L12 | 2023.05.08~ 2024.03.08 | 10 |
|  |  |  |  | (#21): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-0681130 | Female, 43 year old, Taiwan | (#24) (#25) (#26) | - | (#24): Bone Type Implant: Ø4.1× L10 | 2023.05.03~ 2024.01.10 | 8.7 |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0620718 | Female, 49 year old, Taiwan | (#27) | - | (#27): Bone Type Implant: Ø4.8× L10 | 2023.06.06~ 2024.02.05 | 7.30 |
| 2-0821212 | male, 29 year old, Taiwan | (#16) | - | (#16): Bone Type Implant: Ø4.8× L12 | 2022.08.17~ 2023.06.07 | 9.21 |
| 2-540315 | male, 57 year old, Taiwan | (#15) (#17) | - | (#15): Bone Type Implant: Ø4.1× L14 | 2023.02.23~ 2024.05.09 | 14.16 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0550921 | male, 56 year old, Taiwan | (#14) (#15) (#16) | - | (#14): Bone Type Implant: Ø4.1× L10 | 2023.06.26~ 2024.09.30 | 15.4 |
|  |  |  |  | (#15): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0670430 | Female, 45 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.06.026~ 2024.05.29 | 11.3 |
| 2-0520808 | Female, 59 year old, Taiwan | (#12) (#16) (#17) (#21) | - | (#12): Bone Type Implant: Ø3.3× L12 | 2022.11.14~ 2023.06.28 | 7.14 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#21): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-0631113 | male, 48 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.1× L10 | 2022.12.12~ 2024.10.16 | 22.4 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0520701 | Female, 60 year old, Taiwan | - | (#36) (#37) | (#36): Bone Type Implant: Ø4.8× L8 | 2022.07.24~ 2024.03.13 | 19.18 |
|  |  |  |  | (#37): Bone Type Implant: Ø4.8× L8 |  |  |
| 2-0530222 | male, 59 year old, Taiwan | (#14) (#15) (#16) (#17) | - | (#14): Bone Type Implant: Ø4.1× L12 | 2023.07.26~ 2024.04.24 | 8.29 |
|  |  |  |  | (#15): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#16): Bone Type Implant: Ø4.5× L12 |  |  |
| 2-0520722 | male, 60 year old, Taiwan | (#22) (#23) (#25) (#26) | - | (#22): Bone Type Implant: Ø3.3× L12 | 2023.07.26~ 2024.03.13 | 7.16 |
|  |  |  |  | (#23): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0521211 | Female, 59 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.5× L12 | 2023.07.26~ 2024.03.08 | 7.11 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.5× L10 |  |  |
| 2-0391011 | Female, 72 year old, Taiwan | (#15) (#16) (#17) | - | (#15): Bone Type Implant: Ø4.1× L12 | 2023.08.07~ 2024.05.08 | 9.1 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0620601 | male, 50 year old, Taiwan | (#26) | (#46) (#47) | (#26): Bone Type Implant: Ø4.8× L12 | 2023.08.09~ 2024.03.06 | 6.26 |
|  |  |  |  | (#46): Bone Type Implant: Ø4.8× L8 |  |  |
|  |  |  |  | (#47): Bone Type Implant: Ø5.5× L8 |  |  |
| 2-0531104 | male, 58 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.08.09~ 2024.04.03 | 7.25 |
| 2-0540720 | Female, 58 year old, Taiwan | (#25) (#27) | - | (#25): Bone Type Implant: Ø4.1× L10 | 2023.08.21~ 2024.04.01 | 7.11 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-520301-1 | Female, 60 year old, Taiwan | (#11) (#12) (#21) | - | (#11): Bone Type Implant: Ø3.3× L12 | 2023.08.25~ 2024.07.30 | 11.5 |
|  |  |  |  | (#12): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#21): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-0451115 | Female, 66 year old, Taiwan | (#16) (#17) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.08.25~ 2024.04.17 | 7.23 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0470612 | Female, 64 year old, Taiwan | (#12) (#13) | - | (#12): Bone Type Implant: Ø3.3× L12 | 2022.12.28~ 2023.08.25 | 7.28 |
|  |  |  |  | (#13): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-0480323 | Female, 64 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.08.25~ 2024.05.13 | 8.18 |
| 2-0601220 | male, 51 year old, Taiwan | (#24) (#25) (#26) (#27) | - | (#24): Bone Type Implant: Ø4.1× L10 | 2023.09.04~ 2024.06.17 | 9.13 |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L8 |  |  |
| 2-0620911 | Female, 49 year old, Taiwan | (#12) (#13) | - | (#12): Bone Type Implant: Ø3.3× L12 | 2023.09.06~ 2024.04.26 | 7.20 |
|  |  |  |  | (#13): Bone Type Implant: Ø3.3× L14 |  |  |
| 2-530406 | Female, 59 year old, Taiwan | (#27) | - | (#27): Bone Type Implant: Ø4.8× L10 | 2023.09.18~ 2024.03.15 | 5.26 |
| 2-0640409 | male, 48 year old, Taiwan | (#16) (#17) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.09.20~ 2024.06.03 | 8.14 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-380128 | Female, 74 year old, Taiwan | (#26) | (#46) | (#26): Bone Type Implant: Ø4.5× L10 | 2023.10.03~ 2024.04.23 | 6.20 |
|  |  |  |  | (#46): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0540405 | Female, 58 year old, Taiwan | (#27) | - | (#27): Bone Type Implant: Ø4.8× L12 | 2023.10.04~ 2024.05.13 | 7.9 |
| 2-570428-1 | male, 58 year old, Taiwan | - | (#47) | (#47): Bone Type Implant: Ø4.8× L12 | 2023.10.13~ 2024.08.22 | 10.9 |
| 2-0531020 | Female, 59 year old, Taiwan | - | (#46) (#47) | (#46): Bone Type Implant: Ø4.8× L10 | 2023.11.01~ 2024.07.03 | 8.2 |
|  |  |  |  | (#47): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0811005 | male, 30 year old, Taiwan | (#26) |  | (#26): Bone Type Implant: Ø4.8× L10 | 2023.05.08~ 2024.01.12 | 8.4 |
| 2-371211-1 | Female, 74 year old, Taiwan | - | (#46) | (#46): Bone Type Implant: Ø4.8× L10 | 2023.11.02~ 2024.11.02 | 12 |
| 2-600330 | male, 52 year old, Taiwan | - | (#37) | (#37): Bone Type Implant: Ø4.8× L10 | 2023.10.24~ 2024.05.02 | 6.8 |
| 2-0641208 | Female, 47 year old, Taiwan | (#16) (#17) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.11.15~ 2024.09.18 | 10.3 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0451031 | male, 67 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L12 | 2023.02.22 2024.06.19 | 15..28 |
| 2-0380907 | male, 73 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.05.10~ 2024.01.15 | 8.5 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-0660210 | Female, 46 year old, Taiwan | (#16) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.03.13~ 2023.11.17 | 8.4 |
| 2-0741230 | male, 37 year old, Taiwan | (#26) (#27) | - | (#24): Bone Type Implant: Ø3.3× L12 | 2023.04.12~ 2024.01.10 | 8.29 |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-480405 | male, 64 year old, Taiwan | (#21) | - | (#21): Bone Type Implant: Ø4.1× L14 | 2023.11.22~ 2024.04.02 | 4.11 |
| 2-0830718 | male, 29 year old, Taiwan | (#24) (#25) | (#36) | (#24): Bone Type Implant: Ø3.3× L12 | 2023.11.27~ 2024.07.29 | 8.2 |
|  |  |  |  | (#25): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#36): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-0850117 | Female, 27 year old, Taiwan | (#16) | (#45) | (#16): Bone Type Implant: Ø4.5× L10 | 2023.05.10~ 2024.01.12 | 8.2 |
|  |  |  |  | (#45): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-0550112 | male, 57 year old, Taiwan | (#26) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.11.29~ 2024.08.19 | 8.21 |
| 2-300401 | Female, 82 year old, Taiwan | (#12) (#21) | - | (#12): Bone Type Implant: Ø3.3× L12 | 2023.11.30~ 2024.06.06 | 6.7 |
|  |  |  |  | (#21): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-630624 | Female, 49 year old, Taiwan | (#16) (#25) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.12.07 2024.06.28 | 6.21 |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-300406 | Female, 82 year old, Taiwan | - | (#32) (#41) (#42) | (#32): Bone Type Implant: Ø3.3× L12 | 2023.12.11~ 2024.11.15 | 11.4 |
|  |  |  |  | (#41): Bone Type Implant: Ø3.3× L10 |  |  |
|  |  |  |  | (#42): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-0561116 | Female, 56 year old, Taiwan | (#14) (#15) (#16) |  | (#14): Bone Type Implant: Ø4.1× L10 | 2023.12.13~ 2024.08.26 | 8.13 |
|  |  |  |  | (#15): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0490211 | Female, 63 year old, Taiwan | (#24) (#25) (#26) (#27) | - | (#24): Bone Type Implant: Ø4.1× L10 | 2023.12.13 2024.08.19 | 8.6 |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0631002 | male, 48 year old, Taiwan | (#16) (#26) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.04.12~ 2024.03.27 | 11.15 |
| 2-0571220 | male, 55 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.8× L10 | 2023.12.27~ 2024.08.28 | 8.1 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-500824 | Female, 62 year old, Taiwan | (#26) (#27) | - | (#26): Bone Type Implant: Ø4.1× L10 | 2023.12.27 2024.09.02 | 8.6 |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0560420 | Female, 56 year old, Taiwan | (#23) (#25) (#26) (#27) | - | (#23): Bone Type Implant: Ø3.3× L12 | 2024.01.08~ 2024.08.28 | 7.20 |
|  |  |  |  | (#25): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.1× L12 |  |  |
| 2-0710125 | Female, 41 year old, Taiwan | (#13) (#14) (#15) | - | (#13): Bone Type Implant: Ø4.1× L12 | 2024.01.08~ 2024.09.13 | 8.5 |
|  |  |  |  | (#14): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#15): Bone Type Implant: Ø4.1× L10 |  |  |
| 2-0571206-1 | male, 54 year old, Taiwan | (#16) (#17) | - | (#16): Bone Type Implant: Ø4.8× L10 | 2023.05.10~ 2024.01.10 | 8.0 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-511026-1 | male, 61 year old, Taiwan | - | (#43) (#44) (#46) | (#43): Bone Type Implant: Ø4.1× L12 | 2024.01.16~ 2024.08.01 | 6.16 |
|  |  |  |  | (#44): Bone Type Implant: Ø4.1× L12 |  |  |
|  |  |  |  | (#46): Bone Type Implant: Ø4.8× L12 |  |  |
| 2-0590822 | Female, 53 year old, Taiwan | - | (#36) (#37) | (#36): Bone Type Implant: Ø4.8× L10 | 2024.01.22~ 2024.12.16 | 10.24 |
|  |  |  |  | (#37): Bone Type Implant: Ø4.8× L8 |  |  |
| 2-561209-1 | Female, 56 year old, Taiwan | (#25) (#26) (#27) | - | (#25): Bone Type Implant: Ø3.3× L12 | 2024.01.22 2024.09.11 | 7.20 |
|  |  |  |  | (#26): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0428020 | Female, 70 year old, Taiwan | (#11) (#13) (#14) (#16) | - | (#11): Bone Type Implant: Ø3.3× L12 | 2024.01.24 2024.09.25 | 8.1 |
|  |  |  |  | (#13): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#14): Bone Type Implant: Ø3.3× L12 |  |  |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L12 |  |  |
| 2-0460720 | male, 66 year old, Taiwan | (#25) (#26) (#27) | - | (#25): Bone Type Implant: Ø4.1× L12 | 2024.02.19 2024.11.13 | 8.25 |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#27): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0580107 | male, 55 year old, Taiwan | (#17) | - | (#17): Bone Type Implant: Ø4.8× L12 | 2024.02.19 2024.11.06 | 8.18 |
| 2-650112 | Female, 48 year old, Taiwan | (#15) (#17) | - | (#15): Bone Type Implant: Ø4.1× L10 | 2024.02.20 2024.10.11 | 7.21 |
|  |  |  |  | (#17): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-440204 | male, 69 year old, Taiwan |  | (#36) | (#36): Bone Type Implant: Ø4.8× L10 | 2024.02.20 2024.08.13 | 5.24 |
| 2-0520301 | male, 60 year old, Taiwan | (#23) (#24) (#25) (#26) | (#33) (#34) (#35) (#36) | (#23): Bone Type Implant: Ø4.1× L10 | 2024.02.21 2025.01.10 | 10.20 |
|  |  |  |  | (#24): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#25): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#26): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#33): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#34): Bone Type Implant: Ø4.1× L10 |  |  |
|  |  |  |  | (#35): Bone Type Implant: Ø4.8× L10 |  |  |
|  |  |  |  | (#36): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-440803 | Female, 68 year old, Taiwan | (#24) | - | (#24): Bone Type Implant: Ø4.1× L12 | 2024.02.26 2024.09.30 | 6.10 |
| 2-840410 | Female, 28 year old, Taiwan | (#24) (#25) | - | (#24): Bone Type Implant: Ø3.3× L12 | 2024.03.01 2024.09.05 | 6.4 |
|  |  |  |  | (#25): Bone Type Implant: Ø3.3× L12 |  |  |
| 2-430301 | male, 67 year old, Taiwan | - | (#35) | (#35): Bone Type Implant: Ø4.1× L10 | 2024.03.11 2024.08.27 | 5.16 |
| 2-671005 | Female, 45 year old, Taiwan | (#15) (#16) | - | (#15): Bone Type Implant: Ø4.1× L10 | 2024.03.15 2024.10.24 | 7.9 |
|  |  |  |  | (#16): Bone Type Implant: Ø4.8× L10 |  |  |
| 2-0620525 | male, 50 year old, Taiwan | (#25) (#26) | - | (#25): Bone Type Implant: Ø4.1× L10 | 2024.03.18 2024.11.06 | 7.19 |

- List and justify any deviations from the plan:

General method: Post-operative patient (mandibles and maxillae) information is collected to confirm product safety and performance.

The collected number of cases is listed in the table below.

| Follow-up schedule (mandibles and maxillae) |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Years | 2023 | 2024 | 2025 | 2026 | 2027 |
| Estimated Number of Cases | 30 pieces | 30 pieces | 30 pieces | 30 pieces | 30 pieces |
| Actual Number of Cases | 44 | 77 | 183 | - | - |

- Analysis of the findings (whether positive or negative):

- Intended Use

Through intraoral palpation and X-ray imaging (according to Annex A), the bone and gingival growth restoration of 304 cases were able to meet the implant needs and meet the intended use of the device. Table 1 below shows the results.

Table 1. Patient recovery and intended use Statistics

- Side Effects

Table 2 show that, two patients had a foreign body sensation in the mouth, but all the patients’ side effects did not occur.

Evaluation data: After the surgery, the patient will have pain and swelling due to the wound, and after taking the medication prescribed by the doctor, the side effects will improve significantly in 1~2 weeks, and the patient will be free of symptoms by the next visit. The foreign body sensation varies from person to person and may disappear with time. Therefore the device has a low risk of side effects

Table2. Side Effect Statistics

- Survival rate

Among the cases collected so far, molars have the highest number of cases, with #16 and #26 showing significantly higher counts at 39 and 44 cases, respectively. In contrast, data for incisors (#11, #21, #31, and #41) remain sparse, with only a few cases. This further indicates that the primary function of incisors is to bite and tear food, while molars are mainly used for chewing. As a result, molars bear greater force, are used more frequently, and therefore experience more wear. Comparison of Wear between Molars and Incisors: According to the data in the chart, molars are subject to greater biting force and more frequent use, leading to a higher degree of wear. In comparison, incisors are less likely to be damaged under normal usage.

Through Table 4, there are only 3 cases with a survival rate of more than 3 years, 97 cases with a survival rate of more than 1 year, and the rest of the cases are less than 1 year.

Table3. The distribution of tooth position cases

Table4. Case of Follow-up Survival rate

- Summarize the collected data:

Based on the data, there are 97 cases with a survival rate exceeding 12 months, accounting for a significant portion of all samples; however, 204 cases have a survival time of less than 12 months. Notably, there are no cases with survival beyond 60 months, and only 3 cases have survived for over 36 months, indicating that the survival rate of samples is concentrated around 12 months. The data also show that the device has a low risk of side effects, keeping patients in a safe condition and meeting its intended use. Moving forward, patients will be invited for follow-up visits by phone to collect additional clinical data. If fewer than 30 cases are collected within the next year, the PMCF plan will be revised.

## A2 – Screening of scientific literature:

- Summarize all collected clinical data obtained from this activity:

Relevant literature data deemed appropriate. The period covered by the clinical data search is 1983- 2023. The keywords entered for the bibliographic search are:

- Principal Search Terms: Implant, Abutment, Screw, Surgical instrument,Restoration component, Denture, Dental bridge.

- Additional Search Terms (used if required to focus the search): Endosseous, dental, osseointegration, surface, edentulism, edentulous, treatment, therapy.

Results of the Selection Process:

- Sivaswamy, V., & Bahl, V. (2023). Surface Modifications of Commercial Dental Implant Systems: An Overview. J Long Term Eff Med Implants, 33(2), 71-77. doi:10.1615/JLongTermEffMedImplants.2022042612

- Ozkan, A., Çakır, D. A., Tezel, H., Sanajou, S., Yirun, A., Baydar, T., & Erkekoglu, P. (2023). Dental Implants and Implant Coatings: A Focus on Their Toxicity and Safety. J Environ Pathol Toxicol Oncol, 42(2), 31-48. doi:10.1615/JEnvironPatholToxicolOncol.2022043467

- Lemos, C. A. A., Nunes, R. G., Santiago-Júnior, J. F., Marcela de Luna Gomes, J., Oliveira Limirio, J. P. J., Rosa, C. D. D. R. D., . . . Pellizzer, E. P. (2023). Are implant-supported removable partial dentures a suitable treatment for partially edentulous patients? A systematic review and meta-analysis. The Journal of prosthetic dentistry, 129(4), 538-546. doi:10.1016/j.prosdent.2021.06.017

- Inchingolo, A. M., Malcangi, G., Ferrante, L., Del Vecchio, G., Viapiano, F., Inchingolo, A. D., . . . Patano, A. (2023). Surface Coatings of Dental Implants: A Review. Journal of Functional Biomaterials, 14(5). doi:10.3390/jfb14050287

- Cahyaningtyas, N. A., Miranda, A., Metta, P., & Bawono, C. A. (2023). Dental implant macrodesign features in the past 10 years: A systematic review. Journal of Indian Society of Periodontology, 27(2), 131-139. doi:10.4103/jisp.jisp_676_21

- van Oirschot, B., Zhang, Y., Alghamdi, H. S., Cordeiro, J. M., Nagay, B. E., Barao, V. A. R., . . . van den Beucken, J. (2022). Surface Engineering for Dental Implantology: Favoring Tissue Responses Along the Implant. Tissue Eng Part A, 28(11-12), 555-572. doi:10.1089/ten.TEA.2021.0230

- Luke Yeo, I. S. (2022). Dental Implants: Enhancing Biological Response Through Surface Modifications. Dent Clin North Am, 66(4), 627-642. doi:10.1016/j.cden.2022.05.009

- Kim, J. C., Lee, M., & Yeo, I. L. (2022). Three interfaces of the dental implant system and their clinical effects on hard and soft tissues. Mater Horiz, 9(5), 1387-1411. doi:10.1039/d1mh01621k

- Borges, G. A., Barbin, T., Dini, C., Maia, L. C., Magno, M. B., Barão, V. A. R., & Mesquita, M. F. (2022). Patient-reported outcome measures and clinical assessment of implant-supported overdentures and fixed prostheses in mandibular edentulous patients: A systematic review and meta-analysis. The Journal of prosthetic dentistry, 127(4), 565-577. doi:10.1016/j.prosdent.2020.11.005

- Bandiaky, O. N., Lokossou, D. L., Soueidan, A., Le Bars, P., Gueye, M., Mbodj, E. B., & Le Guéhennec, L. (2022). Implant-supported removable partial dentures compared to conventional dentures: A systematic review and meta-analysis of quality of life, patient satisfaction, and biomechanical complications. Clinical and experimental dental research, 8(1), 294-312. doi:10.1002/cre2.521

- Ankita, P., Megha, C., Vashist, S., Ashish, C., Rashi, S., & Jha, D. U. K. (2022). “Dental Implant System: A Systematic Review”. Journal of Pharmaceutical Negative Results, 13, 1076-1083. doi:10.47750/pnr.2022.13.S10.125

- Lafta, H. K., Nada, M. A. E. M., & Rostom, D. A. Soft Tissue Changes Assessment In Screw-Retained Maxillary Implant Supported Complete Denture with Electric Welded Metal Framework Versus Casted One (Rct, Split Mouth Design).

- Vatėnas, I., & Linkevičius, T. (2021). One abutment one time vs. repeatable abutment disconnections in implants, restored with cemented / screw retained fixed partial dentures: Marginal bone level changes. A systematic review and meta-analysis. Stomatologija, 23(2), 35-40. Retrieved from https://www.embase.com/search/results?subaction=viewrecord&id=L636158332&from=export

- Lu, B., Zhang, X., & Liu, B. (2021). A systematic review and meta-analysis on influencing factors of failure of oral implant restoration treatment. Annals of palliative medicine, 10(12), 12664-12677. doi:10.21037/apm-21-3449

- Ferrer, N., Aceituno-Antezana, O., Astudillo-Rozas, W., & Valdivia-Gandur, I. (2021). Genetic Polymorphisms Associated with Early Implant Failure: A Systematic Review. International Journal of Oral and Maxillofacial Implants, 36(2), 219-233. doi:10.11607/jomi.8181

- Bezerra, A. P., Gama, L. T., Pereira, L. J., van der Bilt, A., Peyron, M. A., Rodrigues Garcia, R. C. M., & Gonçalves, T. (2021). Do implant-supported prostheses affect bioavailability of nutrients of complete and partially edentulous patients? A systematic review with meta-analysis. Clin Nutr, 40(5), 3235-3249. doi:10.1016/j.clnu.2021.02.018

- Abdel Baseer, S., Elcharkawi, H., & Dohiem, M. M. (2021). The Effect of Laser Grooved Implant Retained Overdenture on Bone Height Changes. Egyptian Dental Journal, 67(1-January (Fixed Prosthodontics, Removable Prosthodontics and Dental Materials)), 589-596.

- Yussif, N., Rahman, A. A., & Darhous, M. (2020). Evaluation of the health of the peri-implant tissues around immediately placed dental implants in aggressive periodontitis patients versus periodontally healthy individuals in the maxillary esthetic zone: Controlled clinical trial. Journal of Osseointegration, 12(1).

- Knaus, J., Schaffarczyk, D., & Cölfen, H. (2020). On the Future Design of Bio-Inspired Polyetheretherketone Dental Implants. Macromol Biosci, 20(1), e1900239. doi:10.1002/mabi.201900239

- de Avila, E. D., van Oirschot, B. A., & van den Beucken, J. (2020). Biomaterial-based possibilities for managing peri-implantitis. J Periodontal Res, 55(2), 165-173. doi:10.1111/jre.12707

- Abdel-wahab, A., & Tawfik, B. E.-D. A. (2020). Immediate Placement of Dental Implant with Ridge Augmentation Using Autogenous Bone Ring Transplant. Al-Azhar Assiut Dental Journal, 3(1), 33-42.

- Romanos, G. E., Delgado-Ruiz, R., & Sculean, A. (2019). Concepts for prevention of complications in implant therapy. Periodontol 2000, 81(1), 7-17. doi:10.1111/prd.12278

- López-Píriz, R., Cabal, B., Goyos-Ball, L., Fernández, A., Bartolomé, J. F., Moya, J. S., & Torrecillas, R. (2019). Current state-of-the-art and future perspectives of the three main modern implant-dentistry concerns: Aesthetic requirements, mechanical properties, and peri-implantitis prevention. J Biomed Mater Res A, 107(7), 1466-1475. doi:10.1002/jbm.a.36661

- Czumbel, L. M., Kerémi, B., Gede, N., Mikó, A., Tóth, B., Csupor, D., . . . Varga, G. (2019). Sandblasting reduces dental implant failure rate but not marginal bone level loss: A systematic review and meta-analysis. PLoS ONE, 14(5). doi:10.1371/JOURNAL.PONE.0216428

- Almas, K., Smith, S., & Kutkut, A. (2019). What is the Best Micro and Macro Dental Implant Topography? Dent Clin North Am, 63(3), 447-460. doi:10.1016/j.cden.2019.02.010

- Santiago Junior, J. F., Biguetti, C. C., Matsumoto, M. A., Kudo, G. A. H., da Silva, R. B. P., Saraiva, P. P., & Fakhouri, W. D. (2018). Can genetic factors compromise the success of dental implants? A systematic review and meta-analysis. Genes, 9(9). doi:10.3390/genes9090444

- Rupp, F., Liang, L., Geis-Gerstorfer, J., Scheideler, L., & Hüttig, F. (2018). Surface characteristics of dental implants: A review. Dent Mater, 34(1), 40-57. doi:10.1016/j.dental.2017.09.007

- Rasouli, R., Barhoum, A., & Uludag, H. (2018). A review of nanostructured surfaces and materials for dental implants: surface coating, patterning and functionalization for improved performance. Biomater Sci, 6(6), 1312-1338. doi:10.1039/c8bm00021b

- Pellegrini, G., Francetti, L., Barbaro, B., & Del Fabbro, M. (2018). Novel surfaces and osseointegration in implant dentistry. J Investig Clin Dent, 9(4), e12349. doi:10.1111/jicd.12349

- Payne, A. G. T., Alsabeeha, N. H. M., Atieh, M. A., Esposito, M., Ma, S., & Anas El-Wegoud, M. (2018). Interventions for replacing missing teeth: Attachment systems for implant overdentures in edentulous jaws. Cochrane Database of Systematic Reviews, 2018(10). doi:10.1002/14651858.CD008001.pub2

- Anas El-Wegoud, M., Fayyad, A., Kaddah, A., & Nabhan, A. (2018). Bar versus ball attachments for implant-supported overdentures in complete edentulism: A systematic review. Clin Implant Dent Relat Res, 20(2), 243-250. doi:10.1111/cid.12551

- Shibli, J. A., Pires, J. T., Piattelli, A., Iezzi, G., Mangano, C., Mangano, F., . . . Ehrenfest, D. M. D. (2017). Impact of Different Implant Surfaces Topographies on Peri-Implant Tissues: An Update of Current Available Data on Dental Implants Retrieved from Human Jaws. Curr Pharm Biotechnol, 18(1), 76-84. doi:10.2174/1389201017666161221120618

- Reissmann, D. R., Dard, M., Lamprecht, R., Struppek, J., & Heydecke, G. (2017). Oral health-related quality of life in subjects with implant-supported prostheses: A systematic review. J Dent, 65, 22-40. doi:10.1016/j.jdent.2017.08.003

- Nogueira, T. E., Dias, D. R., & Leles, C. R. (2017). Mandibular complete denture versus single-implant overdenture: a systematic review of patient-reported outcomes. J Oral Rehabil, 44(12), 1004-1016. doi:10.1111/joor.12550

- Insua, A., Monje, A., Wang, H. L., & Miron, R. J. (2017). Basis of bone metabolism around dental implants during osseointegration and peri-implant bone loss. Journal of Biomedical Materials Research - Part A, 105(7), 2075-2089. doi:10.1002/jbm.a.36060

- Wilson, T. G., Jr., Miller, R. J., Trushkowsky, R., & Dard, M. (2016). Tapered Implants in Dentistry: Revitalizing Concepts with Technology: A Review. Adv Dent Res, 28(1), 4-9. doi:10.1177/0022034516628868

- Tan, H., Peres, K. G., & Peres, M. A. (2016). Retention of Teeth and Oral Health-Related Quality of Life. J Dent Res, 95(12), 1350-1357. doi:10.1177/0022034516657992

- Smeets, R., Stadlinger, B., Schwarz, F., Beck-Broichsitter, B., Jung, O., Precht, C., . . . Ebker, T. (2016). Impact of Dental Implant Surface Modifications on Osseointegration. Biomed Res Int, 2016, 6285620. doi:10.1155/2016/6285620

- Kern, J. S., Kern, T., Wolfart, S., & Heussen, N. (2016). A systematic review and meta-analysis of removable and fixed implant-supported prostheses in edentulous jaws: post-loading implant loss. Clin Oral Implants Res, 27(2), 174-195. doi:10.1111/clr.12531

- Zohrabian, V. M., Sonick, M., Hwang, D., & Abrahams, J. J. (2015). Dental Implants. Semin Ultrasound CT MR, 36(5), 415-426. doi:10.1053/j.sult.2015.09.002

- Ogle, O. E. (2015). Implant surface material, design, and osseointegration. Dent Clin North Am, 59(2), 505-520. doi:10.1016/j.cden.2014.12.003

- Jemat, A., Ghazali, M. J., Razali, M., & Otsuka, Y. (2015). Surface Modifications and Their Effects on Titanium Dental Implants. Biomed Res Int, 2015, 791725. doi:10.1155/2015/791725

- Dudley, J. (2015). Implants for the ageing population. Aust Dent J, 60 Suppl 1, 28-43. doi:10.1111/adj.12282

- Annunziata, M., & Guida, L. (2015). The Effect of Titanium Surface Modifications on Dental Implant Osseointegration. Front Oral Biol, 17, 62-77. doi:10.1159/000381694

- Albertini, M., Fernandez-Yague, M., Lázaro, P., Herrero-Climent, M., Rios-Santos, J. V., Bullon, P., & Gil, F. J. (2015). Advances in surfaces and osseointegration in implantology. Biomimetic surfaces. Med Oral Patol Oral Cir Bucal, 20(3), e316-325. doi:10.4317/medoral.20353

- Yeo, I. S. (2014). Reality of dental implant surface modification: A short literature review. Open Biomedical Engineering Journal, 8(1), 114-119. doi:10.2174/1874120701408010114

- Branemark, P. I., Adell, R., & Albrektsson, T. (1983). Osseointegrated titanium fixtures in the treatment of edentulousness. Biomaterials, 4(1), 25-28. doi:10.1016/0142-9612(83)90065-0

- Albrektsson, T., Branemark, P. I., & Hansson, H. A. (1983). The interface zone of inorganic implants in vivo: titanium implants in bone. Annals of Biomedical Engineering, 11(1), 1-27. Retrieved from https://www.embase.com/search/results?subaction=viewrecord&id=L13061877&from=export

- List and justify any deviations from the plan:

General method: Apply literature search aligned with the literature search methodology documented in the clinical evaluation report (Doc. # APP I-002_Clinical Evaluation Report of Biomate Implant Systems).

- Confirming the safety of the medical device.

- Confirming the performance of the medical device.

No deviation from Section C2 of PMCF plan.

- Analysis of the findings (whether positive or negative):

| No. | Contribution | Comments |
| --- | --- | --- |
| 1 | SOA | The term “osseointegration” was first introduced by Branemark and colleagues in 1965. Osseointegration of the implant to the bone is affected by the implant stability. Titanium is the current mainstay of the material for dental implants due to its great biocompatibility, corrosion resistance, high strength, low modulus of elasticity, good formability, and machinability. And various methodologies have been developed to modify the surface of titanium implants. The authors investigated the influence of various surface modifications on the surface roughness, bone implant contact, and the success and complication rates of the implants. These are relevant to the design and surface treatment of the subject device. |
| 2 | SOA | The clinical success of dental implantation relies greatly on osseointegration, a process in which an implant and the bone tissue connects to each other structurally and functionally. Surface modifications are performed on the implant to avoid failures and improve osseointegration. Creation of rough implant surface is one major modification method. The goal of creating rough surface is to enhance differentiation, attachment, and growth of bone cells. According to the chemical characteristics, the raw materials used for implants can be divided into following types: metals, ceramics, polymers, and composites. The chemical characteristics, the raw materials used for implants are discussed and are relevant to the state of the art of the subject device. |
| 3 | SOA | Removable partial dentures (RPDs) and implant-supported removable partial dentures (ISRPDs) are common treatments for edentulism. Following indices were reviewed to investigate the performance of clinical outcomes brought by those implants: implant survival rate, marginal bone loss, and patient-reported outcome measures (PROMs). The study concludes: 1) The use of ISRPDs can be a favorable rehabilitation treatment for edentulous patients as they provide high implant survival rates with low values of marginal bone loss; 2) ISRPDs provided more improvement in quality of life and satisfaction than conventional RPDs. Conventional dentures and implant-supported dentures are compared from clinical and patient perspectives. The results show that implant-supported dentures should be considered as a treatment option for edentulous patients. |
| 4 | SOA | Edentulism could result in mental, physical, and social consequences in patients. Dental implantation to restore lost tooth has become one of the common treatment options. And titanium is a frequently used implant materials. The physicochemical properties of the implant outermost layer and its interaction with the surrounding essential tissues play a role in determining the success of osseointegration. A fundamental prerequisite for the long-term success of the implant is biological anchorage between the surface of the dental implant and the bone tissue. Bone response is closely related to the implant surface properties. Hydrophilicity and roughness of the implant surface are two key properties favoring contact of body fluids, clots, and osteoblasts at the implant site, which further lead to better osseointegration. Physio-cellular responses are induced after implantation to osseointegrate the implant to the jawbone. Surface engineering of the dental implant aims to enhance these processes and achieve optimal long-term anchorage of the implant. |
| 5 | SOA | A dental implant is an endosseous material placed into the jaw to support a dental prosthesis. It has been considered an option to treat edentulism due to its high success rate. Osseointegration plays a critical role to the success of dental implantation. And osseointegration affected by several factors including initial stability, which is further affected by the implant macro design. Literature with specific focuses on the implant geometry (shape, length, diameter) as well as thread geometry (pitch, shape, depth) are reviewed. Contributions of the implant geometry and the thread geometry to the initial/primary stability of the implant are discussed. |
| 6 | SOA | The use of dental implants in modern dentistry has shown a staggering increase. Once a dental or orthopedic implant is introduced into bone tissue, immunological responses to a foreign body will be initiated. Since immune response affects its long-term integration, the importance of the initial biological response is reviewed as the starting point for design of novel implant surfaces with improved functionality. The authors concluded the feasibility to modulate immune response with physical surface engineering. Roughening the surface of Titanium (Ti) implants to enhance bone growth is the most used strategy. |
| 7 | SOA | The sandblasted, large-grit, acid-etched (SLA) surface has shown evidence of success in long term clinical investigations using implants made of grade 4 commercially pure titanium (cp-Ti). However, the results of SLA surfaces should be evaluated for implants made of materials other than cp-Ti. The electrochemically oxidized grade 4 cp-Ti has been shown to be successful in long-term clinical investigations. However, other key factors, including implant design, should be considered as additional factors determining success. |
| 8 | SOA | An implant, an abutment, and an artificial crown are the three major components of a dental implant system. While implants are placed into the jawbone, abutments are located at the soft tissue (gingiva) area between the jawbone and the mouth. There are two biological interfaces of the dental implant systems: a hard tissue–implant interface and a soft tissue–abutment interface. The gingiva–abutment interface is also referred as the soft tissue seal. An abutment is divided into three parts and each part is determined by what is attached to it: 1) The cementation part, 2) the transmucosal part, and 3) the connection part. Occlusal force is applied to an abutment through an artificial crown, and then to an implant through the implant–abutment connection. The implant–abutment connection must be sufficiently rigid to endure strong occlusal force. The connection between an abutment and an implant can be simply classified as a butt joint or a friction-fit joint, based on the way these components are assembled. The article investigated the impacts on the success of long-term success and complications with regards to the design of abutment the implant-abutment interface. |
| 9 | SOA | Overdentures and implant-fixed protheses are common approaches to treat edentulism. Both interventions possess cons and pros – overdentures are more cost-effective and less surgically invasive, while implant fixed protheses performed better in occlusal force and required less maintenance. Despite literature comparing the two treatment options are relevantly rich from the clinical perspective, less attention is paid to the patient perspective. It is notable that clinical outcomes might differ from patient perceptions. Patient-reported outcome measures (PROMs) are used to evaluate patient perceptions to overdenture and implant-fixed protheses, and the results are compared. Implant-supported protheses were shown to be favored by patients. And clinical assessment based on the survival rate and marginal bone loss indicated that overdentures were no less efficient than fixed prostheses. |
| 10 | SOA | Distal extension removable partial dentures (DERPDs) are considered an adequate treatment for tooth loss. However, this type of prosthesis is associated with various complications and inconveniences, which altogether leading to patients wearing DERPDs less frequently. Implant-supported removable partial dentures (ISRPDs) have been suggested as a minimally invasive approach for partially edentulous patients and are a suitable alternative to DERPDs. The authors compared ISRPDs and DERPDs with patient reported outcome measures (PROMS) and found: 1) ISRPDs significantly improved patients' quality of life and satisfaction compared to DERPDs, and 2) some mechanical and biological complications were observed following the completion of the ISRPDs. Conventional dentures and implant-supported prostheses were compared and discussed. Results of this review disclosed that there were cons and pros of each treatment option. |
| 11 | SOA | More than 220 implant brands are commercially available worldwide, provided by about 80 manufacturers. The implants differ in the materials, surface treatments, shapes, lengths, and forms. Manufacturers have been seeking for novel surface treatment methodologies to improve the anchorage of the implant in the alveolar bone. The commercially available implants can be categorized into different types: Threaded Implant System, Cylinder Implant System, and Tapered Implant System. To achieve primary stability of the implant, implant design and implant surface treatment play important roles. Commercially available implants are provided with various features. They differ in design, material, and surface treatment, aiming to achieve optimal primary stability of the implant. |
| 12 | S&P | Lafta, H. K. et al. enrolled 11 orally healthy patients aged from 30 to 65 years old. Implants were inserted to the maxilla in the complete edentulous patients. Each patient received six implants. 11 subjects undergone successful implantation. Subjects were followed-up with up to nine months after healing of the implant surgery. No drop-outs observed during the study. No complication / adverse events observed after implantation. Each subject received six implants. In sum, 66 implants were successfully inserted. 100 % success rate of implantation during the test period. Only orally healthy subjects were involved and this is consistent with the indications of the subject device. Only 11 subjects were recruited. The data collected might be biased given the sample size although statistical modifications were applied to correct the bias. |
| 13 | SOA | Disconnections of healing or prosthetic abutment negatively affects the restoration of tooth lost in dentistry. Disconnection and reconnection of the abutment from the implant might disrupt the peri-implant mucosal seal, leading to marginal bone loses. The mis-fit between the abutment framework and the implant, although inevitable during the prosthetic procedures, could generate uneven stresses and strains. This was relevant to complications including screw fracture, framework fracture, and implant loss. Multiple abutment disconnections and reconnections had shown to impact on marginal bone level changes around the implants in patients. Further, multiple (> five times) abutment disconnections and reconnections might influence the soft and hard peri-implant tissue alterations. Not only the success of implant insertion and anchorage, but the connection of abutment to the implant is crucial to the longevity and success of dental implantation. There are multiple factors impacting on the attachment of the abutment to the implant. And disconnection/reconnection might ultimately lead to failure of tooth restoration with implant systems. |
| 14 | SOA | Radiotherapy is correlated to implant failure. Patients who have undergone radiotherapies, whether before or after the dental implantation, exhibited higher likelihood of implant failure comparing to patients who have not been treated with radiotherapy. And peri-implant inflammation results in implant failures as well. It is common that edentulous patients also suffer from periodontitis. Smoking is yet another factor leading to implant failure. Smoking populations are more likely to have failure of oral implants than non-smoking populations. Factors attributing to the failure of the dental implants are discussed in this article, including side-effects / complications. |
| 15 | SOA | Implant failure is linked to not only to the patient’s clinical status but to the genetic factors, such as the genetic polymorphism. Polymorphisms in genes expressing molecules involved in bone inflammation and remodeling processes are suggested to result in early implant failure. The article associated the implant failure with polymorphisms in genes involved in the osseointegration. Though the authors concluded that further investigations were needed, this study provided a possibility for the personalized / customized treatment method of edentulous patients. |
| 16 | SOA | The authors investigated whether implant-supported prostheses are superior to conventional overdentures in terms of masticatory function and nutrient intake. It was found that masticatory improvements for dental fixed implant supported prostheses were better than conventional removable prostheses. Implant-supported prostheses rehabilitation performed better than conventional removable overdentures with regards to masticatory functions. Therefore, it is reasonable to see significant improvements in nutrient intake. |
| 17 | S&P | Abdel Baseer, S. et al. enrolled 10 orally healthy patients aged from 40 to 60 years old. The patient profiles of the two clinical investigations were in alignment with Biomate Medical’s claim of intended patient population. All subjects (10 patients) succeeded in the dental implantation without adverse events. The device performed as claimed by the Biomate Medical. Changes in the crestal bone height were measured. Minor crestal bone loss observed. The dental implantation in all subjects (10 patients) were successful until 12 months post-surgically, without any complications. All implants showed successful osseointegration. Only healthy edentulous subjects enrolled and age range consistent with Biomate Medical’s claim. Limited patient numbers and short follow-up time might not be sufficient to assess the long-term success of implantation and crestal bone loss. |
| 18 | S&P SOA | 70 patients were enrolled in the study by Yussif, N. et al., among whom 35 subjects were orally healthy and 35 subjects were diagnosed with aggressive periodontitis. Note patients with oral infection and inflammation were contraindicated for the use of Biomate Implant Systems, according to the manufacturer. All subjects (70 patients) succeeded in the dental implantation with immediate loading protocol. There was no drop-out during the study period. The device performed as claimed by the Biomate Medical. Six months after the implantation was completed, no peri-implant soft tissue dehiscence, edema, suppuration, pain or mobility was reported in both study groups. All 70 implants were successfully osseointegrated and no failure observed. 100 % success rate of implantation in both study group. The results appeared convincible in terms of the study design, number of patients, and methods of statistical analysis. The subject device performed as intended during the study on the patients in the test and control groups. The results reflected that the Biomate Implant Systems performed well on orally non-healthy patients, such as patients with aggressive periodontitis. |
| 19 | SOA | Over 1300 types of commercial dental implants have been reported in literature. They differ in forms, materials, interface properties, and geometries. Titanium and titanium alloys have been the gold standard as the material for dental implant for decades. However, there are down sides of titanium and its alloys. The disadvantages include hypersensitivity, peri-implant bone overload, high affinity for plaque deposition, and prevalence to corrosion. Despite the negative characteristics of titanium and titanium alloys, they remain the mainstays of the dental implant materials. Other materials like polyetheretherketone (PEEK), on the other hand, have been gaining increasing attentions as the raw material. |
| 20 | SOA | The complications affecting peri-implant tissue are categorized as peri-implant mucositis and peri-implantitis. Peri-implant mucositis is considered the preceding step of the infection process that evokes peri-implantitis, during which a biofilm-induced inflammatory process is initiated. It is clinically diagnosed by bleeding on probing and visual signs of inflammation limited to the soft tissues. The continued inflammation around the implants promotes disease progression of the hard tissues and ultimately leads to peri-implant bone loss. Comparable to periodontal infection, peri-implantitis is a polymicrobial disease, which can also lead to bone resorption and eventually implant loss. Studies have reported infection-related implant loss in 20% of patients during 5-10 years after implant placement. Understanding the pathogenesis causing peri-implantitis is required prior to considering any methods to prevent this disease and to avoid implant failure caused by peri-implant bone loss. |
| 21 | S&P | Abdel-wahab, A., & Tawfik, B. E.-D. A. enrolled 10 patients aged between 24 and 49 years old. The jawbone of all patients in this clinical investigation was severely compromised, which was contraindicated for the subject device as well. Although the demographic information was incomplete, numbers of male and female and ages, it did not limit the validity of the four clinical investigations listed above in terms of the success rate of implantation for up to one year. No complications or adverse events were reported during the study. All implants were successfully osseointegrated without further loading failure. 100 % success rate of implantation. The study showed that the Biomate Implant Systems could be successfully implanted with the immediate auto-bone augmentation. The subject device is not limited to be used on healthy jawbone. |
| 22 | SOA | Marginal crestal bone loss is one complication associated with dental implant. This is caused by surgical trauma, type of implant‐abutment connection, abutment disconnection and reconnection, presence of micro gap, and implant malpositioning. These factors are also referred to as ‘biologic processes associated with bone remodeling. One of the major side-effects of dental implant, marginal crestal bone loss, was discussed in detail. The authors summarized factors related to the marginal crestal bone loss from two perspectives: ‘biologic processes associated with bone remodeling’ and ‘factors associated with bacteria’. |
| 23 | SOA | Marginal bone loss can be caused by accumulation of inflammatory cells at the implant-abutment interface. In addition, micro gaps at the implant-abutment interface can facilitate bacterial colonization. Such inflammation also leads to marginal bone resorption. It can be concluded from literature that preventing formation of the plaque can avoid peri-implant diseases. The main concerns in terms of mechanical properties are related to fatigue. Many studies have shown that the loosening and fatigue fracture of the abutment or prostheses screw are the most common failure modes for screw-retained prostheses, especially in single implant-supported restorations in the premolar and molar regions. Inflammation can also occur at the implant-abutment interface and lead to bone resorption. And the mechanical factor contributing to the loosening of the abutment is fatigue. This would also result in failure of dental rehabilitation. |
| 24 | SOA | Roughness had been suggested to be key to osseointegration and impacts on the success of implantation. Therefore, various methods were developed to create rough surface on the dental implant: sandblasting, acid-etching, anodization, plasma-spraying, coating with different bioactive surfaces and the combination of these. Sandblasting was one of the earliest methods used and had been continuously applied to surface treatment to date. However, it was possible that rough implant surface modified the properties of biofilm formation, leading to the prevalence to bacterial infection. Consequently, the marginal bone around rough implants may be less stable and more vulnerable to peri-implantitis. The authors conducted a review to test the hypothesis that there are significant differences in implant failure rates and marginal bone level changes between sand-blasted and machined dental implants (machined implant can be understood as implants with smooth surface). The review revealed that sandblasted implants are better than machined implants concerning implant failure at 1, 2 and 5–6 years. And no significant differences were found between sandblasted and machined implants after 5 years of follow up, in terms of marginal bone loss. |
| 25 | SOA | The biological fixation between an implant surface and bone is defined as osseointegration (OI). During OI, direct bone-to-implant contact (BIC) is established without an intervening connective tissue layer. The physicochemical properties of the implant surface influence the reaction of surrounding tissue. This occurs between 4 and 14 days after implant insertion. The reactions include upregulation of gene expression and induction of angiogenesis, osteogenesis, and neurogenesis, through cellular responses such as cell adhesion, proliferation, differentiation, and migration. The topographic parameters that have been identified to determine the success of OI include (1) the chemical composition, (2) the hydrophobicity / hydrophilicity of the surface, and (3) the roughness of the implant surface. Factors influencing the success of osseointegration are reviewed and discussed. |
| 26 | SOA | Single nucleotide polymorphisms (SNPs) were suggested to be relevant to osseointegration and the longevity of dental implants. The authors investigated if the genetic factors impacted on the dental implant survival rate. A total of 2130 patients combined from 13 studies were included. No strong evidence that genetic factors can lead to the failure of oral rehabilitation with dental implants were concluded. Aside from the interaction between the bone / tissue and the implant, genetic factors relating to the osseointegration also became of research interest. |
| 27 | SOA | Titanium implant screws are the gold standard for oral implant applications, primarily due to their surpassing biocompatibility and their ability for osseointegration. Different surface modifications have been developed and applied on marketed implants by various subtracting and additive methods, including grit-blasting, acid etching by mineral acids, electrochemical anodic oxidation, calcium-phosphate coatings, or several combinations of these techniques, e.g., combined grit-blasted/acid etched surfaces. Methods of surface modification on the dental implant are developed to increase the wettability, or the hydrophilicity, of the implant. This is to increase the likelihood of successful implantation. |
| 28 | SOA | Physical and chemical surface fabrication and modification techniques commonly used in dental implants include machining, grit blasting (sandblasting), laser treatment, and acid etching combining grit blasting. Other modification techniques include sputtering deposition (SD), plasma spraying (PS)… etc. This article also reviewed about the history of dental implant, including Branemark’s work. It also concluded that infections are the main reason for dental implant failures. Design of the dental implant is also discussed. Each surface modification approach has its cons- and pros, therefore leading to the high variety of commercially available dental implants. |
| 29 | SOA | Implant surface modification by means of laser has been proposed to create a surface with appropriate roughness for enhancing osseointegration. The laser directed onto the implant surface heats and melts the titanium, producing a highly- pure surface, as no foreign material is used to modify or clean the surface (plasma spray, etching, sandblasting), with a fair regularity and a nanostructured topography. This article also discusses the biological reaction after implantation, i.e. explaining the inflammation processes and how osseointegration is achieved. The article highlights the characteristics of laser treatment on the dental implant, which are beneficial to the following osseointegration process. The surface treatment of the subject device echoed the conclusion of this article. |
| 30 | SOA | Edentulism is defined as the state of being without any natural teeth. Complete tooth loss is an irreversible condition. From a prosthodontic perspective, the complete loss of either maxillary or mandibular teeth or both leads to adverse aesthetic and biomechanical sequelae including residual ridge resorption, degenerative changes, impaired masticatory function and loss of neuromuscular control. Interventions used to treat edentulism facilitates improved masticatory function, bite forces, food selection, and quality of life. Interventions discussed in this article include overdenture and removable implant supported prostheses. Different treatment options for edentulism are discussed. Clinical benefits and patient experience are included in this review for conclusion. |
| 31 | SOA | Implant supported overdentures have been proven by many studies to be an effective treatment option for restoring completely edentulous arches, with high success rates. Various attachment systems are actively used to anchor overdentures to the underlying implants, including bar, ball and socket, magnets, and telescopic attachments. Bar and ball attachments are the more commonly employed systems for anchorage. The subject device includes bar and ball abutments which are referred to as the ‘bar and ball attachments’ in this article. The article reviewed and discussed characteristics of these attachment systems in support of prostheses. |
| 32 | SOA | Different surfaces of dental implant are reviewed and discussed: machined surfaces, sandblasted surfaces, plasma sprayed Surfaces, acid-etched surfaces, sandblasted and acid etched surfaces, anodized surfaces, and direct laser metal sintering (DLMS). The article discusses various methods developed up to 2017 in terms of implant surface modification. |
| 33 | SOA | Missing teeth are a common problem among adults of all ages. Not only in European countries but in the US, there are patients suffering from losing ≥1 teeth. Tooth loss, and in particular visible tooth loss (i.e., in the aesthetic zone), is associated with aesthetic problems and reductions in oral health-related quality of life (OHRQoL). Masticatory ability is also influenced by tooth loss. Patients with conventional dentures, particularly wearers of mandibular dentures, may experience looseness or mobility of the denture, differences in salivary flow, oral sensory function and bite force, and an accelerated rate of residual bone resorption. Some patients also experience aesthetic as well as functional problems including difficulties in pronouncing certain sounds and chewing problems. Edentulism as a common problem, various intervention options have been developed for the treatment. This article compared characteristics of common interventions and included patient perspectives. |
| 34 | SOA | Although conventional complete dentures are the most common treatment for edentulous patients worldwide, for older patients in particular, patients experience and complain side-effects of the treatment. The side-effects included impaired function, frequent oral pain and discomfort – especially about the mandibular denture. The authors compared patient reported outcomes for conventional dentures and single implant mandibular overdenture (SIMO). They found marked improvement in satisfaction with the dentures and quality of life measures after SIMO treatment, irrespective of variations in implant treatment protocols and retention systems. |
| 35 | SOA | The authors summarized physio-cellular mechanisms involved during bone remodeling, a critical step to success dental implantation. The force applied to insert the implant, known as the insertion torque, was also reviewed in literature. Adequate implant insertion torque (IT) values (25 – 45 Ncm) have been suggested to prevent micromovements. Implant placement with higher IT (≥ 50 Ncm) has been shown to be more prone to marginal bone loss and recession. In addition, other factors relating to the osseointegration, and peri-implant bone loss were discussed. These included occlusal trauma experiences caused by intraoral disease, factors affecting bone metabolism (e.g., vitamin D), and hyperglycemia. The article discussed not only physio-cellular mechanisms relating to osseointegration, but the insertion torque which is applied to the implant. These factors and processes all play roles in the success of dental implantation. |
| 36 | SOA | Implant design can be characterized by following parameters: implant body shape, thread design, surface modification. Together with the drill protocol used in implant bed preparation, they all directly contribute to the treatment outcome related to bone preservation. A tapered-screw implant design can provide adequate stability because it creates pressure on cortical bone in areas of reduced bone quality. As tapered implant design has evolved over the years, it has become apparent that thread design, placement of the tapered portion, and surface characteristics are all critical parameters when designing a novel tapered implant. All implant variants included in the subject device system are tapered implant. The results of the study supported the design of the subject device is not only state-of-the-art but is safe and effective in teeth rehabilitation. |
| 37 | SOA | Quality of life has been widely recognized as a valid parameter of assessment for a person’s physical and mental health care, including oral health. Measures of oral health–related quality of life (OHRQoL) reflects a true subject-centered approach and represent the views and feelings of people. Tooth loss has a multitude of health-related and socioeconomic repercussions that are associated with 1) oral and systemic diseases; 2) oral health care systems, including the availability and accessibility of dental services; 3) the prevailing philosophies of dental care; 4) health insurance schemes; and 5) other social factors, such as patients’ and dentists’ attitudes and relationships, as well as secular trends and cultural priorities. OHRQoL provides quantitative evaluation from patient perspective of restoring lost teeth. Patient experience can be different from clinical judgement when assessing the success of rehabilitating lost teeth. The article discussed that validity of using OHRQoL to assess patient experience, to gain deeper insights into how edentulism could be treated with regards to how the quality of the patient’s life is changed. |
| 38 | SOA | The central focus of implant development is to minimize bacterial adhesion while promoting the success rate of osseointegration. To guarantee long-term success in clinically challenging conditions, the development of multifunctional surface modifications and coatings is necessary. Up to 2016, surface modification approaches have been of great research interest to enhance the integration of the implant and bone remodeling. |
| 39 | SOA | A complete denture is the classic therapy of full edentulism. However, patients reported insufficient chewing abilities and articulation problems, but also experience psychic strain and social impairment. Studies showed that implant supported prostheses were superior in improving the quality of the patient life after insertion. Nowadays, the insertion and/or restoration of dental implants in edentulous jaws can considered to be one of the basic treatment modalities in a dentist’s everyday practice. The authors reported successful outcome for screw-retained fixed restorations and bar- or ball-retained overdentures in the completely edentulous jaw, including the maxillae and mandible. Furthermore, rough-surfaced implants demonstrated favorable results compared to machined implants. |
| 40 | SOA | A major cause of tooth loss in the United States is periodontitis. Before dental implants, dentures or bridges were used for cosmesis in edentulous patients. A removable denture is a prosthesis supported by the surrounding hard and soft tissues of the oral cavity. A denture may be partial or complete depending on whether a patient is missing some or all of his or her teeth in a single arch. A bridge is a fixed (non-removable) partial denture that is used to replace a missing tooth or several missing teeth. A bridge consists of an artificial tooth or teeth composed of gold, porcelain, or porcelain-fused-to-metal. A crown is used to anchor a bridge to the adjacent normal teeth. Implants have become the state of the art in dentistry and are the most optimal substitute for missing teeth. Alternative treatments for edentulism are discussed, e.g., denture or dental bridge. |
| 41 | SOA | The author discussed the materials used for dental implants, surface treatments, the design of the dental implant, as well as the relationships between osseointegration and surface treatment and implant design. As a primary intervention for lost teeth, dental implant is commercially available and are provided with different characteristics. |
| 42 | SOA | In dental implant, the surface treatment is used to modify the surface topography and surface energy, resulting in an improved wettability, increased cell proliferation and growth, and accelerated osseointegration process. The quality of the dental implant depends greatly on the surface. The author discussed various methods applied in the surface treatment of titanium dental implants. Since the introduction of dental implant to treat edentulism, titanium and its alloys remain the mainstays as the raw material for implant. The article reviewed how surface topography of the titanium implant affect the outcome of osseointegration. Note this article does not discuss implant made of other materials. |
| 43 | SOA | Aged patients do not appear to be associated with different implant success rates compared with younger patients. The quality and quantity of bone available for implant placement are more important predictors of implant success than age. Local bone conditions are both age and site specific. While aging is widely considered to be a factor to affect treatment result, the article concluded on the contrary. It shows that local bone quality plays a more vital role in successful dental implantation. |
| 44 | SOA | The use of titanium dental implants has consistently changed the way of rehabilitating patients in modern dentistry and can count on high long-term survival and success rates. Titanium has been considered the ‘gold standard’ for osseointegration. The authors discussed surface modifications applied to titanium implants: grit blasting, acid etching, combination of blasting and etching, plasma spraying, and anodization. The effects of those surface treatments on osseointegration are also discussed. |
| 45 | SOA | Dental implant osseointegration is a phenomenon that has been studied for a long time. The authors summarized cellular events which characterize osseointegration: protein adsorption, clot formation, granulation tissue formation, provisional matrix formation, interface formation, bone apposition and remodeling. Since the concept of osseointegration is introduced and have been gradually accepted in past decades, researchers studied the cellular mechanisms underlying osseointegration. |
| 46 | SOA | The surface characteristics of dental implants are important to their long-term success. Dental implants made from commercially pure titanium (cp Ti) had shown distinctive longevity for up to 10 years. Yet such implant required longer time to establish. Modifying the surface topography could improve the biocompatibility, provoke osseointegration, and shortening the edentulous period of a patient. The roughness on the implant surface had been believed to improve induces an excellent bone-cell response to the surface. Approaches applied to create rough implant surface included blasting, acid etching, and blasting and acid etching (SLA). To conclude, titanium implants had been the mainstay as the raw material up to 2014. Various approaches had been developed to treat the implant surface, aiming to optimize the time needed for osseointegration. Other raw materials like zirconia, were also in development. |
| 47 | SOA | Focused on titanium implant. The authors evaluated the osseointegrated titanium implants clinically, radiographically, and histologically. The clinical and radiological evaluation indicated that the functioning implants had been surrounded by bone tissues adjacent to the titanium surface. Histological section examination also showed a re-modelled cortical bone and a healthy marrow tissue around the implant, without interposed soft tissues. This is one of the pioneer studies which discussed the result of treating edentulism with titanium implants. The authors pointed out that direct bone to implant contact might not be the sole reason for long-term implant success. The quality of the interface bone was probably of great importance. |
| 48 | SOA | At the microscopic level, surface roughness and porosity provide interlocking of the implant and bone tissue which grows into direct contact with titanium implant. The interface zone includes a tightly adherent titanium oxide layer on the surface of the implant which may be similar to a ceramic material in relation to tissue response. The article also reflects back to Branemark’ s work published in 1977, which showed high success rate of dental implantation in a 10-year follow-up study. Since the concept of osseointegration is introduced, scholars have been dedicated in disclosing how anchorage of the implant into the jaw bone is achieved. |

- Summarize the collected data:

Papers identified in the literature searches and retained following the selection process described in the literature search and review protocol were subject to detailed appraisal using the criteria defined in the literature appraisal protocol. Appraisals and summaries of the published data sets accepted for inclusion in the CER are provided in the table below. The summaries include the significant information and data pertinent to the scientific rationale on safety and/or performance and/or state of the art for the Biomate Implant Systems.

## A3 – PMS publicly listed available databases:

- Summarize all collected clinical data obtained from this activity:

The gathers were conducted from:

- BfArM Field Corrective Actions (Germany)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems with the following limits (unless otherwise specified):

- Category: Medical devices

- Format: Customer information

- Period: Between 2018 and 2025

The following terms were searched:

- Biomate Biomedical

- Biomate Implant Systems

- Biomate System

- Osstem Implant

- TS III Implant

- Geass

- Way Implant System

| Search | BfArM Reference | Manufacturer/Product | Event Description or Device Problem or Rationale for Exclusion |
| --- | --- | --- | --- |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | 11 results |  |  |
|  | 21994/20 | Medtronic GmbH/HeartWare™ Ventricular Assist Device (HVAD™) System | Excluded – Not relevant to the subject device or similar devices. |
|  | 11195/23 | Abiomed, Inc./Impella 5.5® with SmartAssist® heart pump | Excluded – Not relevant to the subject device or similar devices. |
|  | 07731/22 | Medtronic GmbH/Percept™ PC Implantable Neurostimulator (INS) Model B35200 | Excluded – Not relevant to the subject device or similar devices. |
|  | 27900/21 | Medtronic GmbH/AbreTM venous self-expanding stent system | Excluded – Not relevant to the subject device or similar devices. |
|  | 11309/21 | SynCardia Systems, LLC/The SynCardia TAH-t System | Excluded – Not relevant to the subject device or similar devices. |
|  | 11055/21 | Boston Scientific International S.A./INGENIO™ Family DR EL Pacemakers and CRT-Ps (92705305-FA) And Hydrogen Induced Accelerated Battery Depletion (92289212-FA) | Excluded – Not relevant to the subject device or similar devices. |
|  | 20857/20 | Boston Scientific International S.A./EMBLEM™ Subcutaneous Implantable Cardioverter Defibrillators (S-ICDs) (Models A209 and A219) EMBLEM™ S-ICD Subcutaneous Electrode (Model 3501) | Excluded – Not relevant to the subject device or similar devices. |
|  | 12786/18 | Jarvik Heart, Inc./JHI-212 Y Cable, JHI-202 Lithium-ion Battery Cable, JHI-201 Pedestal Cable, JHI-204 Extension Cable | Excluded – Not relevant to the subject device or similar devices. |
|  | 03769/18 | Boston Scientific International S.A./INGEVITY™ MRI active fixation pacing lead manual | Excluded – Not relevant to the subject device or similar devices. |
|  | 03938/18 | implantcast GmbH/MUTARS® HD Coupling C-O-M | Excluded – Not relevant to the subject device or similar devices. |
|  | 12241/17 | implantcast GmbH/Trochanter plate for MUTARS® Proximal Femur Revision | Excluded – Not relevant to the subject device or similar devices. |
|  | 03193/25 | Medtronic GmbH/Pipeline™ Vantage Embolization Device | Excluded – Not relevant to the subject device or similar devices. |

- MHRA Medical Device Alerts (UK)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems with the following limits (unless otherwise specified):

- Alert type: Field safety notice

- Issued after: 01/08/2019 (DD/MM/YYYY)

- Issued before: 01/08/2025 (DD/MM/YYYY)

- Medical specialty: Dentistry

The following terms were searched:

- Biomate Biomedical

- Biomate Implant Systems

- Biomate System

- Osstem Implant

- TS III Implant

- Geass

- Way Implant System

| Search | MHRA Reference | Manufacturer/Product | Event Description or Device Problem or Rationale for Exclusion |
| --- | --- | --- | --- |
|  | No results | N/A | N/A |
|  | 1 messages |  |  |
|  | Field Safety Notices: 21 to 25 November 2022 | Various (9 field safety notices, 2 corrective actions) | Excluded – Not relevant to the subject device. |
|  | 1 messages |  |  |
|  | Field Safety Notices: 21 to 25 November 2022 | Various (9 field safety notices, 2 corrective actions) | Excluded – Duplicated. |
|  | 1 messages |  |  |
|  | Field Safety Notices: 21 to 25 November 2022 | Various (9 field safety notices, 2 corrective actions) | Excluded – Duplicated. |
|  | 3 messages |  |  |
|  | Field Safety Notices: 21 to 25 November 2022 | Various (9 field safety notices, 2 corrective actions) | Excluded – Duplicated. |
|  | No results |  |  |
|  | 1 messages |  |  |
|  | Field Safety Notices: 21 to 25 November 2022 | Various (9 field safety notices, 2 corrective actions) | Excluded – Duplicated. |

- SwissMedic Recalls & FSCA (Switzerland)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems. The following limitations were applied (unless otherwise specified):

- From 01/08/2019 (DD/MM/YYYY)

- To 01/08/2025 (DD/MM/YYYY)

The following terms were searched:

- Biomate Biomedical

- Biomate Implant Systems

- Biomate System

- Osstem Implant

- TS III Implant

- Geass

- Way Implant System

| Search | SwissMedic Ref. | Manufacturer/Product | Event Description or Device Problem or Rationale for Exclusion |
| --- | --- | --- | --- |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | Vk_20200525_06 | GROUPE SEBBIN SAS/Mammary Implants | Excluded – Not relevant to the subject device or similar devices. |
|  | No results | N/A | N/A |
|  | 6 results |  |  |
|  | Vk_20230421_11 | Abiomed Inc./Impella 5.5 SmartAssist Heart Pump | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20220412_14 | Medtronic Inc./CareLink SmartSync Device Manager | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20220331_16 | Medtronic Inc./Percept PC BrainSense | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20180911_13 | Cardiac Pacemaker Inc/ACCOLADE, PROPONENT, ESSENTIO, and ALTRUA 2 pacemakers and VISIONIST and VALITUDE cardiac resynchronization therapy pacemakers (CRT-Ps) | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20210603_04 | Cardiac Pacemaker Inc/INGENIO, VITALIO, and ADVANTIO pacemakers and INLIVEN, INTUA, and INVIVE cardiac resynchronization therapy pacemakers (CRT-Ps) | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20200602_01 | OCON Medical Ltd./ IUB Ballerine Intrauterine Device | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20250130_24 | Micro Therapeutics Inc/ Pipeline Vantage Embolization Device | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20220331_16 | Medtronic Inc./ Percept PC BrainSense | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20240515_33 | Medtronic Inc./ Percept, Activa SC, DBS Clinician Programmer Application, DBS Pocket Adaptor | Excluded – Not relevant to the subject device or similar devices. |
|  | Vk_20211013_07 | Medtronic Inc./ Percept PC BrainSense | Excluded – Not relevant to the subject device or similar devices. |

- TGA DAEN (Australia)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems.

The following terms were searched:

- Biomate Biomedical

- Biomate Implant Systems

- Biomate System

- Osstem Implant

- TS III Implant

- Geass

- Way Implant System

Note: the events reported from July 1st, 2012 to Aug 1st, 2025. No more recent data are available.

| Search | DAEN Reference | Manufacturer/Product | Event Description or Device Problem or Rationale for Exclusion |
| --- | --- | --- | --- |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | 4 results |  |  |
|  | 45449 | Osstem Implant Co Ltd/Osstem Implant Co. Ltd - Osstem TS System CA Fixture | Loss of stability. Event type: Implantable device failure. |
|  | 45544 | Osstem Implant Co Ltd/Osstem Implant Co.,Ltd - Osstem TS System CA Fixture | Fixture stability was too low to performing prosthetic procedure. Event type: Implantable device failure. |
|  | 45547 | Osstem Implant Co Ltd/ Osstem Implant Co.,Ltd - Osstem TS System CA Fixture - Screw endosteal dental implant, two-piece | Due to the low stability implant has been explanted. Event type: Implantable device failure. |
|  | 45546 | Osstem Implant Co Ltd/ Osstem Implant Co.,Ltd - Osstem TS System SA Fixture - Screw endosteal dental implant, two-piece | Due to the low stability, implants were explanted. Event type: Implantable device failure. |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |
|  | No results | N/A | N/A |

- US FDA MAUDE Database (USA)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems with the following limitations (unless otherwise specified):

- Date Range: from 01/08/2019 to 01/08/2025 (DD/MM/YYYY)

The following terms were searched:

- Advanced Search: Manufacturer – Biomate Biomedical

- Advanced Search: Brand Name – Biomate Implant Systems

- Advanced Search: Brand Name – Biomate System

- Advanced Search: Manufacturer – Osstem

- Advanced Search: Brand Name – TS III Implant

- Advanced Search: Manufacturer – Geass

- Advanced Search: Brand Name – Way Implant System

| Search | Report Number | Product Description | Event Type | Event Description or Device Problem or Rationale for Exclusion |
| --- | --- | --- | --- | --- |
|  | No results | N/A | N/A | N/A |
|  | No results | N/A | N/A | N/A |
|  | No results | N/A | N/A | N/A |
|  | 12 results |  |  |  |
|  | 3003394081-2023-00004 | OSSTEM IMPLANT CO., LTD. TSIII SA; DENTAL IMPLANT | Injury | Dental implant fracture. |
|  | 3003394081-2023-00002 | OSSTEM IMPLANT CO., LTD. TSIII SA IMPLANT; DENTAL IMPLANT | Injury | Dental implant fracture. |
|  | 3003394081-2023-00001 | OSSTEM IMPLANT CO., LTD. MS SA IMPLANT; DENTAL IMPLANT | Injury | Dental implant fracture. |
|  | 3003394081-2023-00003 | OSSTEM IMPLANT CO., LTD. SSIII SA IMPLANT; DENTAL IMPLANT | Injury | Dental implant fracture. |
|  | 3003394081-2022-00006 | OSSTEM Co. LTD./ USII SA DENTAL IMPLANT | Injury | Dental implant failed to osseointegrate. |
|  | 3003394081-2022-00007 | OSSTEM Co. LTD./TS III DENTAL IMPLANT | Injury | Dental implant failed to osseointegrate. |
|  | 3003394081-2022-00005 | OSSTEM Co. LTD./MS SA DENTAL IMPLANT | Injury | Dental implant fractured. |
|  | 3003394081-2022-00004 | OSSTEM CO., LTD./TSIII SA DENTAL IMPLANT | Injury | Dental implant fractured. |
|  | 3003394081-2022-00003 | OSSTEM CO., LTD./TSIII SA DENTAL IMPLANT | Injury | Dental implant fractured. |
|  | 3003394081-2022-00002 | OSSTEM CO., LTD./TSIII SA DENTAL IMPLANT | Injury | Dental implant fractured. |
|  | 3003394081-2022-00001 | OSSTEM CO., LTD./TSIII SA DENTAL IMPLANT | Injury | Dental implant fractured. |
|  | 3003394081-2021-00001 | OSSTEM CO., LTD. /OSSTEM IMPLANT SYSTEM DENTAL IMPLANT | Malfunction | Cover screw would not disengage from the dental implant. |
|  | 3003394081-2025-00001 | OSSTEM CO., LTD./ TS III SA | Injury | Dental implant facture. |
|  | 3003394081-2025-00002 | OSSTEM CO., LTD./ TS III SA | Injury | Dental implant facture. |
|  | No results | N/A | N/A | N/A |
|  | No results | N/A | N/A | N/A |
|  | No results | N/A | N/A | N/A |

- US FDA Medical Device Recalls (USA)

Date(s) of search: Aug 1st, 2025

The following searches were conducted to capture the reportable events related to the Biomate Implant Systems with the following limitations (unless otherwise specified):

- Date Range: from 08/01/2019 to 08/01/2025 (MM/DD/YYYY)

The following terms were searched:

- Advanced Search: Recalling firm – Biomate Biomedical

- Advanced Search: Product – Biomate Implant Systems

- Advanced Search: Product – Biomate System

- Advanced Search: Recalling firm – Osstem

- Advanced Search: Product – TS III Implant

- Advanced Search: Recalling firm – Geass

- Advanced Search: Product – Way Implant System

| Search | Recall Number | Recalling Firm | Product | Reason for Recall or Rationale for Exclusion |
| --- | --- | --- | --- | --- |
| 1. | No results | - | - | - |
| 2. | No results | - | - | - |
| 3. | No results | - | - | - |
| 4. | No results | - | - | - |
| 5 | No results | - | - | - |
| 6 | No results | - | - | - |
| 7 | No results | - | - | - |

- List and justify any deviations from the plan:

General method: Implement updated PMS database search per PMS database search according to the PMS database search performed in the clinical evaluation report (Doc. # APP I-002_Clinical Evaluation Report of Biomate Implant Systems).

- Identifying previously unknown side-effects (related to the procedures or to the medical devices).

- Monitoring the identified side-effects and contraindications

- Identifying and analyzing emergent risks

- Ensuring the continued acceptability of the benefit-risk ratio

- Identifying the need for medical device reporting and FSCA

- Identify the need for vigilance reporting and FSCA.

- Identifying possible systematic misuse or off-label use of the device

- Identify the need for trend reporting

No deviation from Section C3 of PMCF plan.

- Analysis of the findings (whether positive or negative):

The following table: Results of Reportable Events Database Searches

| Database | Relevant Results for Subject Devices | Relevant Results for Similar Devices |
| --- | --- | --- |
| BfArM Field Corrective Actions (Germany) | 0 | 0 |
| MHRA Medical Device Alerts (UK) | 0 | 0 |
| SwissMedic Recalls & FSCA (Switzerland) | 0 FSCA | 0 FSCA |
| TGA DAEN (Australia) | 0 | 4 |
| US FDA MAUDE Database (USA) | 0 | 14 |
| US FDA Medical Recalls (USA) | 0 FSCA | 0 FSCA |

Summary of the results are provided in the following sections.

- BfArM Field Corrective Actions (Germany):

No events related to the subject and similar devices were identified in this database.

- MHRA Medical Device Alerts (UK):

No adverse events related to the subject and similar devices were identified in this database.

- SwissMedic Recalls & FSCA (Switzerland) :

No adverse events related to the subject and similar devices were identified in this database.

- TGA DAEN (Australia) :

Implant stability is a core clinical factor to the success rate of dental implant. Low implant stability or loss of implant stability would impact on the following prosthetic procedure. In other words, implant stability plays a critical role in the performance of dental implant systems. However, the underlying reasons for the reported events derived from the search results are unclear. Root-cause analysis of the adverse events would be required for elucidation.

- US FDA MAUDE Database (USA) :

Amongst the fourteen included for discussion, two injury events resulted from failure to osseointegrate. This reflected the link between successful osseointegration and endosseous dental implantation. The other twelve events resulted from implant fracture. However, without narratives provided by the manufacturer, further investigations would be needed to identify the root-cause of implant fracture and how implant fractures led to injuries and malfunctions.

- US FDA Medical Device Recalls (USA) :

No adverse events related to the subject and similar devices were identified in this database.

- Summarize the collected data:

From Reportable Event Databases

Public PMS databases have been reviewed on  Aug 1st, 2025.  The following reportable events and field safety corrective actions (FSCAs)/recalls with clinical impact were identified:

- 0 events related to the Biomate Implant Systems have been reported;

- 18 events related to similar devices to the Biomate Implant Systems were reported (1 malfunctions, 13 injuries, 0 deaths and 4 FSCAs).

The events and FSCAs were reported in TGA DAEN and US FDA MAUDE databases though UK, Switzerland, Australian and German databases were also consulted.

The reportable event database review identified safety events with clinical impact related to the Biomate Implant Systems or similar devices that are presented below:

- Loss of stability. [Osstem Implant Co Ltd/Osstem Implant Co. Ltd - Osstem TS System CA Fixture]

- Fixture stability was too low to perform prosthetic procedure. [Osstem Implant Co Ltd/Osstem Implant Co.,Ltd - Osstem TS System CA Fixture]

- Due to the low stability implant has been explanted. [Osstem Implant Co Ltd/ Osstem Implant Co.,Ltd - Osstem TS System CA Fixture - Screw endosteal dental implant, two-piece]

- Dental implant failed to osseointegrate. [OSSTEM Co. LTD./ USII SA DENTAL IMPLANT]

- Dental implant fractured. [OSSTEM Co. LTD./MS SA DENTAL IMPLANT]

- Cover screw would not disengage from the dental implant. [OSSTEM CO., LTD. /OSSTEM IMPLANT SYSTEM DENTAL IMPLANT]

## B – Evaluation of clinical data relating to similar devices:

Biomate has elected not to use the clinical data from an equivalent (clinical, technical and biological characteristics) device. In the event, there is a device considered equivalent, its data will be considered as a similar device.

| Product name of similar device | Results discussed | References used to get the results (publications, part of technical documentation from this similar device) |
| --- | --- | --- |
| TS III System | N/A | N/A |
| Way implant system | N/A | N/A |

## C – Impact of the results on the technical documentation:

| Clinical Evaluation Report (CER): | APP I-002_Clinical Evaluation Report |
| --- | --- |
| Relevant information analysed and monitored: The above data will be discussed in the following sections of the Clinical Evaluation Report (APP I-002): Appendix D Clinical Literature Data Appraisal and Summary Appendix E Summary of Reportable Events Database Data Appendix G Summary of Clinical Data Generated and Held by the Manufacturer |  |
| Risk Management File (RMF): | APP G-004_Implant System Risk Management Report |
| Relevant information analysed and monitored: Reference to BI-TD-0010_ Sections 1. Conclusions of the benefit-risk determination. |  |

## D – Reference to any applicable common specification(s), harmonized standard(s) or guidance document(s) applied:

Common specification(s) to comply with, if applicable:

N/A

Harmonised standards to apply, if applicable:

N/A

Guidance on PMCF, if applicable:

MDCG 2020-7 Post-market clinical follow-up (PMCF) Plan Template: A guide for manufacturers and notified bodies (April 2020)

- Statistical data of unqualified quality

- Unqualified quality data

| year/month | Raw material return rate (%) < 2% (Month) | Process unqualified rate (%) < 3% (Month) | Unqualified rate of finished products (%) | Finished product scrap rate (market) (%) | Finished product scrap rate (in-factory) (%) |
| --- | --- | --- | --- | --- | --- |
| 2025.01 | 0% | 0% | 0% | 0% | 0% |
| 2025.02 | 0% | 0% | 0% | 0% | 0% |
| 2025.03 | 0% | 0% | 0% | 0% | 0% |
| 2025.04 | 0% | 0% | 0% | 0% | 0% |
| 2025.05 | 0% | 0% | 0% | 0% | 0% |
| 2025.06 | 0% | 0% | 0% | 0% | 0% |
| 2025.07 | 0% | 0% | 0% | 0% | 0% |
| 2025.08 | 0% | 0% | 0% | 0% | 0% |
| 2025.09 | 0% | 0% | 0% | 0% | 0% |
| 2025.10 | 0% | 0% | 0% | 0% | 0% |
| 2025.11 | 0% | 0% | 0% | 0% | 0% |
| 2025.12 | 0% | 0% | 0% | 0% | 0% |

- Summarize the collected data of unqualified quality:

Based on past analyses and data from early 2024, the raw material return rate for Biomate Implant Systems had previously exceeded the set standard of 5%. Specifically, in January 2024, the raw material return rate reached as high as 57.46%, was 16.94% in February, and remained at 0.16% even in March and April. Concurrently, the in-process unqualified rate for January and March 2024 were 1.42% and 1.55% respectively. These unqualified quality issues primarily stemmed from packaging materials, such as the plastic bottles for implants and labels. In response to the issues identified above, the company has implemented corrective actions, including returning unqualified raw materials and processing some specially accepted goods. Notably, according to the latest data from January to December 2025, all unqualified quality indicators, including raw material return rate, in-process unqualified rate, finished product unqualified rate, finished product scrap rate (market), and finished product scrap rate (in-factory), have all decreased to 0%. This clearly indicates that the corrective actions taken previously have effectively resolved the related issues and led to a significant improvement in the overall quality status. These unqualified issues did not affect product safety, and all necessary corrective actions have been completed. Therefore, the company evaluates this risk as acceptable.

- Sales and Complaint Data

- Sales data

Data Attribute Definition: The sales volume data presented in this section represents Units Placed on the Market (devices distributed to distributors or end-users) during the reporting period.

Sales Volume by Basic UDI-DI and Region: In accordance with MDCG 2022-21 Annex IV, the following table summarizes the sales volume stratified by Basic UDI-DI, EMDN code, and geographical region (EEA+TR+XI vs. Non-EEA). Detailed sales data by individual catalogue number is provided in Annex B.

Table 5-1: Sales Volume (Units Placed on the Market)

| Basic UDI-DI | Device Family / EMDN Code | Region | 2025 year | 2024 year | 2023 year | 2022 year | Total |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 471987540Implant5K | Implant /P01020101 | EEA+TR+XI | 85 | 1 | 1,473 | 631 | 2,190 |
|  |  | Worldwide | 8,304 | 7,685 | 3,522 | 8,435 | 27,946 |
| 471987540AbutmentVC | Abutment /P01020180 | EEA+TR+XI | 0 | 49 | 2,105 | 367 | 2,521 |
|  |  | Worldwide | 12,091 | 10,980 | 6,426 | 11,282 | 40,779 |
| 471987540ScrewIIbUN | Screw (IIb) /P01020180 | EEA+TR+XI | 85 | 1 | 1,473 | 631 | 2,190 |
|  |  | Worldwide | 8,715 | 8,089 | 3,944 | 9,488 | 30,236 |
| 471987540Surgical(IIa)FP | Surgical Instruments / Z12110102 | EEA+TR+XI | 0 | 0 | 88 | 215 | 303 |
|  |  | Worldwide | 332 | 600 | 1,217 | 1,756 | 3,905 |

Note 1: "EEA+TR+XI" includes EU Member States, Turkey, and Northern Ireland.

Note 2: Non-EEA represents sales to the rest of the world (e.g., Taiwan, Asia).

Note 3: EMDN codes are derived from the Declaration of Conformity (Doc #: BI-IM-S001).

- Complaint Data

The complaint data presented in this report was collected from 2017 to 2024. During this period, all customer feedback and adverse events were recorded and assessed.

To ensure compliance with MDR requirements and facilitate global harmonization, all complaint records have been retrospectively classified using the IMDRF Adverse Event Terminology (AET).

• Total Complaints: A total of 28 complaints were received during the reporting period.

• Overall Complaint Rate: The overall complaint rate is 0.02% based on the total sales volume.

For a detailed breakdown of failure modes, trend analysis, and side-by-side comparison by year, please refer to Section 5.4. Detailed individual records are listed in Annex C.

- Rational and description of any preventive and corrective actions taken:

Refer to Annex C: Preventive and corrective actions list for the detailed description of each event, root cause investigation, and the specific corrective actions (CAPA) implemented.

- Summarize the collected data of the complaint:

From 2017 to 2024, all collected complaint data has been retrospectively classified using IMDRF Adverse Event Terminology (AET) to facilitate standardized trend analysis. As detailed in Annex C (Comprehensive Analysis of Complaints with IMDRF Codes), the primary device problems identified were categorized by IMDRF Annex A codes (e.g., A010201 Failure to Osseointegrate, A040101 Fracture).

To satisfy the requirement for side-by-side comparison and trend assessment, the following table summarizes the complaint occurrence rates per product family and specific failure mode. The rate (%) is calculated based on the sales volume (units distributed) for the respective year.

Table 5.4-1: Statistical Trend Analysis of Complaints by IMDRF Code

| Basic UDI-DI / Device Model | IMDRF AET Annex A Code & Term | EEA + TR + XI |  | Worldwide |  | Trend Assessment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | N | Rate % | N | Rate % |  |
| 471987540Surgical(I)XE /3AA-034, 3AA-039 | A040502 - Corroded |  |  |  |  |  |
|  | Year (2017) | 0 | 0 % | 1 | 0.00 % | No recurrence since 2017 |
|  | Year (2018) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2019) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2020) | 0 | 0 % | 0 | 0 % |  |
| 471987540RestorationKR / BSMURT000000A | A1202 - Decoupling |  |  |  |  |  |
|  | Year (2021) | 0 | 0 % | 1 | 0.00 % | No recurrence |
|  | Year (2022) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540AbutmentVC / 4AA-D17 | A020102 - Nonstandard Device |  |  |  |  |  |
|  | Year (2021) | 0 | 0 % | 1 | 0.00 % | No recurrence |
|  | Year (2022) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540AbutmentVC / 4AA-D19 | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 1 | 0.01 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(IIa)FP / 3AA-N24 | A040101 - Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 5 | 0.28 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(IIa)FP / 3AA-N06 | A020101 - Dull, Blunt |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 3 | 0.17 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540RestorationKR / 6AA-061 | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 1 | 0.76 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(I)XE / BSSITR000000A | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 1 | 6.25 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 2 | 25.00 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(I)XE / BSSITR000000A | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(I)XE / 3AK-D03 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(I)XE / 3AK-056 | A040101 - Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 5 | 0.00 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(IIa)FP / 3AA-016, 3AA-014, 3AA-015 | A020101 - Dull, Blunt |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 3 | 0.17 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540AbutmentVC / 4AA-E01 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2023) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540RestorationKR / 6AA-085 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Surgical(I)XE / 3AA-056 | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 2 | 12.50 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540AbutmentVC / 4AA-I05 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Implant5K / 1AA-003, 1AA-004, 1AA-022 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 0 | 0 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Implant5K / 1AA-015, 1AA-017, 1AA-018, 1AA-025 | A010201 - Failure to Osseointegrate |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 4 | 0.11 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540Implant5K / 1AA-005, 1AA-022 | A010402 - Migration |  |  |  |  |  |
|  | Year (2023) | 0 | 0 % | 1 | 0.03 % | No recurrence |
|  | Year (2024) | 0 | 0 % | 0 | 0 % |  |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |
| 471987540AbutmentVC / 4AA Series | A040101 – Fracture |  |  |  |  |  |
|  | Year (2024) | 0 | 0 % | 0 | 0 % | Decreasing |
|  | Year (2025) | 0 | 0 % | 0 | 0 % |  |

* Note: (N) = Total number of events. Rate (%) calculation is based on Sales Volume (units distributed) for the respective year.

Analysis of Trends and Conclusion:

Based on the statistical analysis in Table 5.4-1 and the detailed investigation in Annex C, the following conclusions are drawn regarding the safety and performance of the device:

- Overall Trend: The majority of device problems identified in 2022 and 2023 (e.g., A020101 Dullness, A010103 Shape issues) show a "Decreasing" or "No recurrence" trend in the current reporting year (2024). This indicates that the corrective actions (CAPAs) implemented, such as manufacturing parameter adjustments and design revisions, have been effective.

- Specific Focus - Osseointegration (A010201): An increase was observed in 2023 (Rate 0.03%). However, data for 2024 shows 0 cases, confirming that this was not a systemic product defect but likely associated with isolated clinical factors. The trend is assessed as Decreasing.

- New Issue Monitoring: A minor fracture issue (A040101) in the Abutment family was noted in 2024 (Rate 0.01%). This is classified as a "New Issue (Under Monitoring)". The rate is within the acceptable risk level defined in the Risk Management File.

- Health Impact (Annex F): According to the impact analysis, these events resulted in No Health Consequences (F26) or required minor Device Revision (F1905). No serious injuries were reported.

In conclusion, no statistically significant increasing trends that would alter the benefit-risk profile were identified. We continue to monitor these risks through the PMS system.

- Reportable Events Trending

Internal Biomate company records of reportable events were examined; No reportable events were reported from 2015 to 2025. In addition, internal Biomate company records of FSCAs without a reportable event were assessed as part of reportable events trending.  A summary of reportable events and FSCAs without a reportable event is shown in follow table.

| Device Model | Year | Sales Data | Number of Events Reported to Authorities |
| --- | --- | --- | --- |
| Bone Type Implant family | 2025 | 21,018 | 0% (0 case) |
|  | 2024 | 19,733 | 0% (0 case) |
|  | 2023 | 14,915 | 0% (0 case) |
|  | 2022 | 23,395 | 0% (0 case) |
|  | 2021 | 14,943 | 0% (0 case) |
|  | 2020 | 19,027 | 0% (0 case) |
|  | 2019 | 21,880 | 0% (0 case) |
|  | 2018 | 18,324 | 0% (0 case) |
|  | 2017 | 17,767 | 0% (0 case) |
|  | 2016 | 9,452 | 0% (0 case) |
|  | 2015 | 5,604 | 0% (0 case) |
| BioSmart Implant family | 2025 | 260 | 0% (0 case) |
|  | 2024 | 550 | 0% (0 case) |
|  | 2023 | 331 | 0% (0 case) |
|  | 2022 | 405 | 0% (0 case) |
|  | 2021 | 465 | 0% (0 case) |
|  | 2020 | 727 | 0% (0 case) |
|  | 2019 | 102 | 0% (0 case) |
|  | 2018 | 0 | 0% (0 case) |
|  | 2017 | 0 | 0% (0 case) |
|  | 2016 | 0 | 0% (0 case) |
|  | 2015 | 0 | 0% (0 case) |
|  | TOTAL | 188,898 | 0% (0 case) |

- Conclusions of the PSUR

- Summary of Safety and Performance Findings

- Device Failure (Screw Fracture): In 2024, one complaint regarding a broken screw was received. Upon investigation, it was determined that the specification chosen by the doctor was inappropriate, causing excessive stress on the screw and leading to its fracture. Therefore, it is recommended to provide the doctor with suitable specifications to prevent recurrence.

- Supply Chain Integrity: In 2024, no complaints were received regarding product damage caused by environmental factors (e.g., high or low temperatures) or handling (e.g., hits or falls) during storage or transportation. This confirms that current logistic controls are effective.

- State of the Art (Literature Review): The articles identified during the literature search (1983–2023) were thoroughly evaluated. The data confirms that the safety and performance of the Biomate implant system remain consistent with modern technologies.

- Update Strategy: Typically, literature collection for dental clinical evaluations does not need to be conducted annually due to relatively slow technological changes; a 5-year cycle is considered feasible and efficient. However, if significant technological changes, regulatory updates, or new clinical evidence arise, the literature update will be conducted in advance to ensure timeliness.

- Clinical Experience (Implants): In 2025, we completed the collection of 183 clinical experience cases. Additionally, within the timeframe from March 7, 2016, to February 1, 2025, we gathered a total of 304 cases.

- Survival Analysis: 97 cases exhibited a survival rate exceeding 12 months. Conversely, 204 cases had a survival time of less than 12 months. Notably, there were no cases with survival beyond 60 months, and only 3 cases survived for over 36 months, indicating that the current data is concentrated around the 12-month mark.

- Safety Conclusion: The data shows that the device has a low risk of side effects, ensuring patient safety and confirming its intended use.

  - Justification for Waiving Clinical Investigations for Surgical Instruments (Addressing NC 10)

Based on the data presented in this PSUR, the "4-Pillar Evidence Strategy" confirms that residual risks for Surgical Instruments are acceptable without specific clinical investigations:

Pillar 1: PMS Objective Evidence

- Analysis of sales and complaint records (2021–2024) confirms market stability and performance.

- Out of 3,990 Class IIa units distributed (as detailed in Section 5.4), only 4 performance-related complaints were recorded, resulting in a negligible failure rate of 0.10%.

Pillar 2: Physical Mechanism Analysis

- Citing Delgado-Ruiz et al. (2018) to demonstrate that drill wear and metallic particle release are known inherent physical characteristics of surgical drills.

- These risks are effectively mitigated through adherence to standardized irrigation and usage protocols defined in the manufacturer’s instructions.

Pillar 3: Biological Safety Profile

- Citing Wilson et al. and Zhurakivska et al. to confirm that the biological response to metallic debris is a well-characterized and acceptable risk.

- The use of established medical-grade material (SS420J2) ensures that no additional animal or human testing is required to confirm biocompatibility.

Pillar 4: Clinical Performance Benchmarking

- Citing Aldhohrah et al. and Sutariya et al. to provide high-level clinical evidence that standardized surgical instrumentation achieves high implant survival rates (91.2%–100%).

- These data validate that the current surgical instrumentation meets the state-of-the-art requirements for clinical safety and performance.

Conclusion：

Based on the proven low failure rate (0.10%) and established scientific consensus, the residual risks are deemed acceptable without the necessity for additional clinical investigations.

The data collected will be fed back into the APP G-004 Risk Assessment Report and the APP I-002 Clinical Evaluation Report.

Based on above description, we will revise the PMS plan and PMCF plan.