Table 1. Comprehensive Analysis of Medical Device Complaints with IMDRF Codes (Annex A, C, F)

| Single Number | Date of Occurrence / Country | IMDRF Annex A Code | IMDRF Annex A Term | Complaint Description | MDRF Annex F Code | IMDRF Annex C Code | Action Taken/ Status & Effectivenes |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 170206001 | 2017/02/06 Taiwan | A040502 | Corroded | The doctor reported that our Torque Ratchet (3AA-034), Stopper (No information provided), Screw driver (No information provided) and Implant driver (3AA-039) have rusted condition. | F26 (No Health Consequences or Impact) | C0601 (Degradation Problem Identified) | Inform the doctor that the drill has a service life and if surface rust is found, please do not use it again. Status: Closed Effectiveness: Resolved. No recurrence of corrosion reported for this part since 2017. |
| 210223001 | 2021/02/23 Taiwan | A1202 | Decoupling | The physician responded that the Multi-Unit Abutment Adapter (BSMURT000000A) slipped with a torque of less than 30N on the clinical bed. | F26 (No Health Consequences or Impact) | C16 (Manufacturing/Production Problem Identified) | 1. Quality assurance has finished screening the finished products and transfers the qualified products to the production control and the unqualified products to the isolation area. 2. Make hexagonal plug gauges as acceptance standard. Status: Closed Effectiveness: Effective. Zero slipping incidents reported in subsequent batches. |
| 210315001 | 2021/03/15 Taiwan | A020102 | Nonstandard Device | 15°Angled Abutment D4.0H8.5G1 /SD-Hex (4AA-D17) is too tight and the outer hex size is too large to assemble the implant. This product was found to be abnormal when the technician was working on it, and it was not sent to the clinic yet, so the company replaced a 15° Angled Abutment to the technician. | F2601 - Problem identified before clinical use/exposure | C16 -Manufacturing/ Production Problem Identified | 1. Revise the size of the interfering external hexagon. 2. Check and record once an hour before the production process.3. Quality assurance classify Angled Abutment as special product control. Status: Closed Effectiveness: Effective. Dimensional accuracy verified; no recurrence. |
| 220114001 | 2022/01/14 Jordan | A010103 | Inadequacy of Device Shape and/or Size | - Agent complains 15° Angled Abutment (GH 1) of old version cannot be assembles with Lab Analog. - Although 15° Angled Abutment (GH 1) of new version can be assembles with Lab Analog, it is so tight the GH2. Thus quality of product shall be improved assembling exactly. | F26 - No Health Consequences or Impact | C0402 - Incompatible Component/ Accessory | 1. After 15° Angled Abutment (GH 1) of old version has applied for exchange goods, we send the goods to Sales Department in January 14, 2022. 2. After we analyze the problem, the product has no issue by assembling, please forward the analysis to the agent. Status: Closed. Effectiveness: Resolved. Customer accepted the explanation and replacement. |
| 220331001 | 2022/03/31 Taiwan | A040101 | Fracture | We confirm the description. Because patient’s Opening-mouth-wide is not enough, drill cannot move up-down. Thus drill is damage when it stuck with steal-ring. | F26 - No Health Consequences or Impact | C0706 - Stress Problem Identified | 1. When BIOnavi Kit had launched, it has no related feedback for BIOnavi Initial Drill L5. We will consider adopting L5 into BIOnavi Kit. 2. If customer needs the spec of L5, they can buy individual. Status: Closed. Effectiveness: Resolved. Solution provided for specific patient anatomy cases. |
| 220331002 | 2022/03/31 Taiwan | A020101 | Dull, Blunt | Shape of BIOnavi Final Drill is designed with Arc groove and straight edge, in order to it can collect the bone of patient and it can prevent laser mark position rusty. Therefore, because the designed is weak for cutting, drilling is difficult. | F26 - No Health Consequences or Impact | C060202 - Improper Physical Structure | Since BIOnavi has feedback. We had revised the design in January, 2022 and doctor's feedback is well, after doctor test in training. We are producing the product in May, 2022 and we unify exchanging BIOnavi Final Drill of new version. When we launch the product, we forward the information to doctor. Status: Closed. Effectiveness: Effective. Positive feedback received on new design; cutting efficiency restored. |
| 220622001 | 2022/06/22 Taiwan | A010103 | Inadequacy of Device Shape and/or Size | Patient’s tooth-wide is small, so Close-Tray stuck the adjacent tooth and it cannot be fixed position. Furthermore, RD-Implant is no issue of the stuck, because Back-tooth-wide is wide than front-tooth-wide. | F26 - No Health Consequences or Impact | C19 - No Device Problem Found | If customer needs D4.0 of Close-Tray, we suggest custom-made for the doctor by Aiosen. Status: Closed Effectiveness: Resolved. Case-specific solution provided. |
| 220622002 | 2022/06/22 Taiwan | A010103 | Inadequacy of Device Shape and/or Size | BioSmart Torque Ratchet can stir normally, but up-down move is floating easily. | F26 - No Health Consequences or Impact | C060202 - Improper Physical Structure | The original function is valid, so the complaint is not established. Status: Closed Effectiveness: Confirmed. Product functions within specifications. |
| 220726001 | 2022/07/26 Taiwan | A040101 | Fracture | - We doubt about that abutment screw is reused, because abutment screw has stretching phenomenon. Thus, abutment screw is fracture. - BioSmart Torque Ratchet has two versions. The torque value of New versions is about 5N higher than old versions. Thus, when doctor use a torque value of 30N, it is actually 35N. | F26 - No Health Consequences or Impact | C23 - Usage Problem Identified | We have a training session for doctor how to identify new or old of Torque Ratchet and doctor can exchange new version of Torque Ratchet to old version. Status: Closed Effectiveness: Effective. User education provided to prevent misuse. |
| 220801001 | 2022/08/01 Taiwan | A040101 | Fracture | Original factory reply: Trephine cannot aim at and catch the Implant when Trephine is the high-speed. Furthermore Trephine crashes the Implant, so the Implant is fracture. | F1903 - Device Explantation | C23 - Usage Problem Identified | 1. Original factory recommend: After using low-speed aim at and catch the Implant, use a high-speed. 2. We write the Instruction of Trephine into the manual. Status: Closed Effectiveness: Effective. IFU updated to clarify procedure and reduce risk. |
| 220912001 | 2022/09/12 Taiwan | A040101 | Fracture | Doctor utilizes long-pin version of Implant Driver. Because long-pin could be shaken easily, front-pin is fracture. | F26 - No Health Consequences or Impact | C23 - Usage Problem Identified | 1. We have a short-pin version of Implant Driver and after we provide the product to doctor, doctors' feedback is well. 2. If doctor happen fracture of long-pin version, doctor can exchange the goods. Status: Closed Effectiveness: Effective. Short-pin version resolved the stability issue for users. |
| 220913001 | 2022/09/13 Taiwan | A040101 | Fracture | Doctor utilizes long-pin version of Implant Driver. Because long-pin could be shaken easily, front-pin is fracture. | F26 - No Health Consequences or Impact | C23 - Usage Problem Identified | 1. We have a short-pin version of Implant Driver and after we provide the product to doctor, doctors' feedback is well. 2. If doctor happen fracture of long-pin version, doctor can exchange the goods. Status: Closed Effectiveness: Effective. Short-pin version resolved the stability issue for users. |
| 220913002 | 2022/09/13 Taiwan | A040101 | Fracture | Doctor utilizes long-pin version of Implant Driver. Because long-pin could be shaken easily, front-pin is fracture. | F26 - No Health Consequences or Impact | C23 - Usage Problem Identified | 1. We have a short-pin version of Implant Driver and after we provide the product to doctor, doctors' feedback is well. 2. If doctor happen fracture of long-pin version, doctor can exchange the goods. Status: Closed Effectiveness: Effective. Short-pin version resolved the stability issue for users. |
| 220920001 | 2022/09/20 Taiwan | A020101 | Dull, Blunt | We analyze about that parameter of Electro-polish passivation set overtime, so ability of cutting reduces. Therefore, doctor need to push hard drilling. | F26 - No Health Consequences or Impact | C16 -Manufacturing/ Production Problem Identified | 1. Original factory reply: They inform the problem to MF department and improve ability of cutting. 2. If customer has old version of Counter Sink, they can exchange the goods. Status: Closed Effectiveness: Effective. Process parameters adjusted; cutting performance restored. |
| 221024001 | 2022/10/24 Taiwan | A040101 | Fracture | We analyze about that Hex-part of Abutment Screw is too long and Abutment Screw has over-part of Abutment when Angled abutment assembles with Implant. Thus, dental technician need to cut the over-part, so Abutment Screw is lost ability of fixing. | F2601 - Problem identified before clinical use/exposure | C16 -Manufacturing/ Production Problem Identified | 1. We design a new version of Abutment Screw that Hex-part is short. If customer has old version of Abutment Screw they can exchange the goods. 2. Because the spec of Angled Abutment 25° has less demand, we don't exchange the old version of Abutment Screw. When BiomatSmart launch a long time stably in the market, we unity exchanging the old version of Abutment Screw to new. Status: Closed Effectiveness: Effective. Design improvement eliminated the need for manual modification. |
| 221103001 | 2022/11/03 Taiwan | A020101 | Dull, Blunt | We analyze about that parameter of Electro-polish passivation set overtime, so ability of cutting reduces. Therefore, doctor need to push hard drilling. | F26 - No Health Consequences or Impact | C16 -Manufacturing/ Production Problem Identified | 1. Original factory reply: They inform the problem to MF department and improve ability of cutting. 2. If customer has old version of Counter Sink, they can exchange the goods. Status: Closed Effectiveness: Effective. Process parameters adjusted; cutting performance restored. |
| 230109001 | 2023/01/09 Thailand | A040101 | Fracture | There are two incidents that Abutment Screw is fracture when doctor twist Abutment Screw by hands. | F26 - No Health Consequences or Impact | C19 - No Device Problem Found | Prof. Dr. Tao Chiang think about that torque force is 10Ncm by utilizing hands, so it is impossible to twist the fracture of Abutment Screw. After we have taken a video that Abutment Screw could not be twisted to the fracture by utilizing hands, we send the video to agent. Status: Closed Effectiveness: Resolved. Agent/Customer accepted the evidence provided. |
| 230201001 | 2023/02/01 Taiwan | A010103 | Inadequacy of Device Shape and/or Size | Utilizing short-pin version of Implant Driver can pick up the spec of RD 4.8 or 5.5 Implant, but they have no hand-feeling that Implant Driver assembles with the Implant. | F26 - No Health Consequences or Impact | C0602 - Inappropriate Material (Structure) | Old version of Implant' hole is a little big than new version. Due to there are not another same incident, so we discuss about that does not sell old version to CLOSER company in the future. Status: Closed Effectiveness: Resolved. Sales restriction prevents recurrence with this client. |
| 230505001 | 2023/05/05 Taiwan | A010103 | Inadequacy of Device Shape and/or Size | Implant Driver (3AA-056) cannot pick up the Implant and cannot pick up any spec of Implant in Demo Box. | F2601 - Problem identified before clinical use/exposure | C16 - Manufacturing/ Production Problem Identified | 1. After Implant Driver changes C-Buckle, it can ready to use. 2. Suggesting Sales and Marketing Department exchange that product of some have problem of assembly in Demo Box. 3. If any products have unusual problem subsequently, please recall the product and send the product to R&D and QA to analyze. Status: Closed Effectiveness: Effective. Assembly issue corrected; monitoring for recurrence. |
| 230705001 | 2023/07/05 Thailand | A040101 | Fracture | Positioner Abutment (4AA-I05) is fracture, because chewing force of patient is too strong. | F1905 - Device Revision or Replacement | C0706 - Stress Problem Identified | Please, sale staff forward to re-do Implant to doctor according to Prof. Dr. Tao Chiang. QA is going to continue the case of follow-up. Status: Closed Effectiveness: Resolved. Patient treated; monitoring for trend in chewing force related failures. |
| 230706002 | 2023/07/06 Thailand | A040101 | Fracture | Doctor made a Full-Mouth-Reconstructionutilize of Upper-Jaw by two Implants. When the patient returns visit, doctor discovered that screw of Positioner Abutment is fracture. Thus, doctor re-do Positioner Abutment for patient. After, the patient returns visit again, doctor discovered that Implant is a fracture. | F1905 - Device Revision or Replacement | C20- No Findings Available | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; insufficient evidence to determine root cause. |
| 230706003 | 2023/07/06 Thailand | A010201 | Failure to Osseointegrate | Doctor cannot push the Implant in for first time, so doctor utilize Tap and lock-force is over than 50 N-cm to complete the operation. After four weeks, the Implant is failure. The patient recovered in four months. Doctor completed the operation of Implant in Sep. 4, 2022. | F1905 - Device Revision or Replacement | C23 - Usage Problem Identified) | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; failure linked to initial excessive torque. |
| 230706004 | 2023/07/06 Thailand | A010201 | Failure to Osseointegrate | Doctor cannot implant the product by Guide-Plate. When doctor pick up the Guide-Plate, the Implant pick up together. | F1905 - Device Revision or Replacement | C0501 - Inadequate Labelling and/or Instructions for Use | Marketing staff confirm to miss the Instructions of specific product. Marketing staff revise the Instruction completely and sale staff forward the information to doctor. Agent and doctor agree. Status: Closed Effectiveness: Effective. Instructions updated and communicated; no similar complaints since. |
| 230706005 | 2023/07/06 Thailand | A010402 | Migration | Patient is in healthy condition, but the Implant is failure to Osseointegrate. | F1905 - Device Revision or Replacement | C20- No Findings Available | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; cause of migration remains undetermined. |
| 230706006 | 2023/07/06 Thailand | A010201 | Failure to Osseointegrate | Patient has D4-Bone. After the product implants to loosen, doctor chooses other spec of Implant. | F1905 - Device Revision or Replacement | C19 - No Device Problem Found | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; attributed to patient bone quality. |
| 230706007 | 2023/07/06 Thailand | A010201 | Failure to Osseointegrate | The Implant is failure to Osseointegrate. | F1905 - Device Revision or Replacement | C20- No Findings Available | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; insufficient info to identify device fault. |
| 230706008 | 2023/07/06 Rumania | A040101 | Fracture | The Implant is fracture. | F1905 - Device Revision or Replacement | C20- No Findings Available | After the doctor re-did the dental implant surgery, the patient has recovered their health. Status: Closed Effectiveness: Resolved. Patient recovered; insufficient info to identify device fault. |
| 240528001 | 2024/05/28 Taiwan | A040101 | Fracture | The screw of Biomate-Plus broke. | F1905 - Device Revision or Replacement | C23- Usage Problem Identified | After analysis, it was determined that the specification chosen by the doctor was not appropriate, which caused excessive stress on the screw, leading to the fracture. Therefore, it is recommended to provide the doctor with suitable specifications. Status: Closed Effectiveness: Resolved. Appropriate specifications recommended to prevent recurrence. |

Table 2. IMDRF - Statistical Trend Analysis Table

| Basic UDI-DI / Device Model | IMDRF AET Annex A Code & Term | EEA + TR + XI |  | Worldwide |  | Trend Assessment |
| --- | --- | --- | --- | --- | --- | --- |
|  |  | N | Rate % | N | Rate % |  |
| 471987540Surgical(I)XE /3AA-034, 3AA-039 | A040502 - Corroded |  |  |  |  |  |
|  | Year (2017) | 0 | 0.00% | 1 | < 0.01% | No recurrence since 2017 |
|  | Year (2018) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2019) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2020) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540AbutmentVC / 4AA-D19 | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 0.01% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(IIa)FP / 3AA-N24 | A040101 - Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 0.06% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(IIa)FP / 3AA-N06 | A020101 - Dull, Blunt |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 0.06% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(I)XE / BSSITR000000A | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 6.25% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(I)XE / BSSITR000000A | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 6.25% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(I)XE / 3AK-D03 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | < 0.01% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(I)XE / 3AK-056 | A040101 - Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 3 | < 0.01% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(IIa)FP / 3AA-016, 3AA-014, 3AA-015 | A020101 - Dull, Blunt |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 2 | 0.11% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540AbutmentVC / 4AA-E01 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2022) | 0 | 0.00% | 1 | 0.01% | No recurrence |
|  | Year (2023) | 0 | 0.00% | 0 | 0.00% |  |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540RestorationKR / 6AA-085 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 0 | 0.00% | 1 | 1.92% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Surgical(I)XE / 3AA-056 | A010103 - Inadequacy of Device Shape and/or Size |  |  |  |  |  |
|  | Year (2023) | 0 | 0.00% | 2 | 12.50% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540AbutmentVC / 4AA-I05 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 0 | 0.00% | 1 | 0.02% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Implant5K / 1AA-003, 1AA-004, 1AA-022 | A040101 – Fracture |  |  |  |  |  |
|  | Year (2023) | 1 | 0.07% | 2 | 0.06% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Implant5K / 1AA-015, 1AA-017, 1AA-018, 1AA-025 | A010201 - Failure to Osseointegrate |  |  |  |  |  |
|  | Year (2023) | 0 | 0.00% | 4 | 0.11% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540Implant5K / 1AA-005, 1AA-022 | A010402 - Migration |  |  |  |  |  |
|  | Year (2023) | 0 | 0.00% | 1 | 0.03% | Decreasing |
|  | Year (2024) | 0 | 0.00% | 0 | 0.00% |  |
| 471987540AbutmentVC / 4AA Series | A040101 – Fracture |  |  |  |  |  |
|  | Year (2024) | 0 | 0.00% | 1 | 0.01% | New Issue (Under Monitoring) |

* (N): Total number

* Rate (%): calculation is based on Sales Volume (units distributed) for the respective year.